## WaveCalUV paper
No. | Input quantity | Estimated uncertainty |  | Probability distribution | Coverage factor | Standard uncertainty |  | Sensitivity coefficient |  | Uncertainty contribution |  | Output quantity | Measurement model | Measurement result
i | Xi | ue(xi) | unit |  | k | u(xi) | unit |  | unit | ui(y) | unit | Y | Y = F(Xi) | Y ± u(y) for a confidence level of 95 %
1 | Theoretical value spectral line in vacuum | 0.00017321 | nm | Normal | 1 | 0.00017321 | nm | 1 |  | 0.00017321 | nm
2 | Modified Edlen equation calculation | 7.3905e-08 | 1 | Normal | 2 | 3.6953e-08 | 1 | -3286.7815 | nm/1 | 0.00012146 | nm
3 | Laboratory temperature effect | 4 | °C | Triangular | √6 | 1.633 | °C | 0.00035275 | nm/°C | 0.00057604 | nm
4 | Laboratory humidity effect | 30 | %RH | U-shaped | √2 | 21.2132 | %RH | 4.5e-06 | nm/%RH | 9.5459e-05 | nm
5 | Laboratory air pressure effect | 4 | kPa | Rectangular | √3 | 2.3094 | kPa | 0.0011722 | nm/kPa | 0.0027072 | nm
6 | Wavelength calibration equation offset | 286.5196 | steps | Normal | 1 | 286.5196 | steps | 0.00034717 | nm/steps | 0.09947 | nm
7 | Wavelength reproducibility | 0.005 | nm | Rectangular | √3 | 0.0028868 | nm | 1 |  | 0.0028868 | nm
8 | Mechanical resolution of monochromator | 0.025 | nm | Rectangular | √3 | 0.014434 | nm | 1 |  | 0.014434 | nm
9 | Resolution used in calibration | 0.1 | nm | Rectangular | √3 | 0.057735 | nm | 1 |  | 0.057735 | nm
10 | Repeatability | 83.7086 | steps | Normal | 1 | 83.7086 | steps | 0.00034717 | nm/steps | 0.029061 | nm
[9 empty] | uc(y) | 0.11957 | nm
[9 empty] | U (k=2) | 0.3 | nm
6 | Steep-side method offset |  |  | Normal | 1 | 0 |  |  |  | 0 | nm

## WaveCalUV
No. | Input quantity | Symbol | Value |  | Estimated uncertainty |  | Probability distribution | Coverage factor | Standard uncertainty |  | Sensitivity coefficient |  |  | Uncertainty contribution |  | Output quantity | Measurement model | Measurement result
i | Xi |  | xi | unit | ue(xi) | unit |  | k | u(xi) | unit | ci |  | unit | ui(y) | unit | Y | Y = F(Xi) | Y ± u(y) for a confidence level of 95 %
1 | Theoretical value spectral line in vacuum | Usl | 253,728 296,815 435,956 | nm | 0.00017321 | nm | Normal | 1 | 0.00017321 | nm | csl | 1 |  | 0.00017321 | nm
2 | Modified Edlen equation calculation | UEd | 1.0002 | 1 | 7.3905e-08 | 1 | Normal | 2 | 3.6953e-08 | 1 | cEd | -3286.7815 | nm/1 | 0.00012146 | nm
3 | Laboratory temperature effect | UT | 23.3 | °C | 4 | °C | Triangular | √6 | 1.633 | °C | cT | 0.00035275 | nm/°C | 0.00057604 | nm
4 | Laboratory humidity effect | URH | 52.5 | %RH | 30 | %RH | U-shaped | √2 | 21.2132 | %RH | cRH | 4.5e-06 | nm/%RH | 9.5459e-05 | nm
5 | Laboratory air pressure effect | Uap | 86.773 | kPa | 4 | kPa | Rectangular | √3 | 2.3094 | kPa | cap | 0.0011722 | nm/kPa | 0.0027072 | nm
6 | Wavelength calibration equation offset | Uo | 286.5196 | steps | 286.5196 | steps | Normal | 1 | 286.5196 | steps | co | 0.00034717 | nm/steps | 0.09947 | nm
7 | Wavelength reproducibility | Uλ | 0.005 | nm | 0.005 | nm | Rectangular | √3 | 0.0028868 | nm | cλ | 1 |  | 0.0028868 | nm
8 | Mechanical resolution of monochromator | UMres | 0.05 | nm | 0.025 | nm | Rectangular | √3 | 0.014434 | nm | cMres | 1 |  | 0.014434 | nm
9 | Resolution used in calibration | UCres | 0.2 | nm | 0.1 | nm | Rectangular | √3 | 0.057735 | nm | cCres | 1 |  | 0.057735 | nm
10 | Repeatability | Ur | 83.7086 | steps | 83.7086 | steps | Normal | 1 | 83.7086 | steps | cr | 0.00034717 | nm/steps | 0.029061 | nm
[13 empty] | uc(y) | 0.11957 | nm
[13 empty] | U (k=2) | 0.3 | nm
6 | Steep-side method offset | Ussmo |  |  |  |  | Normal | 1 | 0 |  | cT |  |  | 0 | nm

## WaveCalVisIR paper
No. | Input quantity | Estimated uncertainty |  | Probability distribution | Coverage factor | Standard uncertainty |  | Sensitivity coefficient |  | Uncertainty contribution |  | Output quantity | Measurement model | Measurement result
i | Xi | ue(xi) | unit |  | k | u(xi) | unit | ci | unit | ui(y) | unit | Y | Y = F(Xi) | Y ± u(y) for a confidence level of 95 %
1 | Theoretical value spectral line in vacuum | 0.0001 | nm | Normal | 1 | 0.0001 | nm | 1 |  | 0.0001 | nm
2 | Modified Edlen equation calculation | 4.1991e-08 | 1 | Normal | 2 | 2.0995e-08 | 1 | 0.99977 | nm/1 | 2.0991e-08 | nm
3 | Laboratory temperature effect | 4 | °C | Triangular | √6 | 1.633 | °C | 0.000681 | nm/°C | 0.0011121 | nm
4 | Laboratory humidity effect | 30 | %RH | U-shaped | √2 | 21.2132 | %RH | 9.2857e-06 | nm/%RH | 0.00019698 | nm
5 | Laboratory air pressure effect | 4 | kPa | Rectangular | √3 | 2.3094 | kPa | 0.0022569 | nm/kPa | 0.005212 | nm
6 | Steep-side method offset | 36 | steps | Normal | 1 | 36 | steps | 0.00069431 | nm/steps | 0.024995 | nm
7 | Wavelength calibration equation offset | 72 | steps | Normal | 1 | 72 | steps | 0.00069431 | nm/steps | 0.04999 | nm
8 | Wavelength reproducibility | 0.005 | nm | Rectangular | √3 | 0.0028868 | nm | 1 |  | 0.0028868 | nm
9 | Mechanical resolution of monochromator | 0.05 | nm | Rectangular | √3 | 0.028868 | nm | 1 |  | 0.028868 | nm
10 | Resolution used in calibration | 0.05 | nm | Rectangular | √3 | 0.028868 | nm | 1 |  | 0.028868 | nm
11 | Repeatability | 15.5563 | steps | Normal | 1 | 15.5563 | steps | 0.00069431 | nm/steps | 0.010801 | nm
[9 empty] | uc(y) | 0.070313 | nm
[9 empty] | U (k=2) | 0.2 | nm

## WaveCalVisIR
No. | Input quantity | Symbol | Value |  | Estimated uncertainty |  | Probability distribution | Coverage factor | Standard uncertainty |  | Sensitivity coefficient |  |  | Uncertainty contribution |  | Output quantity | Measurement model | Measurement result
i | Xi |  | xi | unit | ue(xi) | unit |  | k | u(xi) | unit | ci |  | unit | ui(y) | unit | Y | Y = F(Xi) | Y ± u(y) for a confidence level of 95 %
1 | Theoretical value spectral line in vacuum | Usl | 696,735 912,547 966,044 | nm | 0.0001 | nm | Normal | 1 | 0.0001 | nm | csl | 1 |  | 0.0001 | nm
2 | Modified Edlen equation calculation | UEd | 1.0002 | 1 | 4.1991e-08 | 1 | Normal | 2 | 2.0995e-08 | 1 | cEd | 0.99977 | nm/1 | 2.0991e-08 | nm
3 | Laboratory temperature effect | UT | 24.4 | °C | 4 | °C | Triangular | √6 | 1.633 | °C | cT | 0.000681 | nm/°C | 0.0011121 | nm
4 | Laboratory humidity effect | URH | 42.1 | %RH | 30 | %RH | U-shaped | √2 | 21.2132 | %RH | cRH | 9.2857e-06 | nm/%RH | 0.00019698 | nm
5 | Laboratory air pressure effect | Uap | 86.525 | kPa | 4 | kPa | Rectangular | √3 | 2.3094 | kPa | cap | 0.0022569 | nm/kPa | 0.005212 | nm
6 | Steep-side method offset | Ussmo | 36 | steps | 36 | steps | Normal | 1 | 36 | steps | cT | 0.00069431 | nm/steps | 0.024995 | nm
7 | Wavelength calibration equation offset | Uo | 72 | steps | 72 | steps | Normal | 1 | 72 | steps | co | 0.00069431 | nm/steps | 0.04999 | nm
8 | Wavelength reproducibility | Uλ | 0.005 | nm | 0.005 | nm | Rectangular | √3 | 0.0028868 | nm | cλ | 1 |  | 0.0028868 | nm
9 | Mechanical resolution of monochromator | UMres | 0.1 | nm | 0.05 | nm | Rectangular | √3 | 0.028868 | nm | cMres | 1 |  | 0.028868 | nm
10 | Resolution used in calibration | UCres | 0.1 | nm | 0.05 | nm | Rectangular | √3 | 0.028868 | nm | cCres | 1 |  | 0.028868 | nm
11 | Repeatability | Ur | 15.5563 | steps | 15.5563 | steps | Normal | 1 | 15.5563 | steps | cr | 0.00069431 | nm/steps | 0.010801 | nm
[13 empty] | uc(y) | 0.070313 | nm
[13 empty] | U (k=2) | 0.2 | nm

## UcTable
Term | Symbol | Examples
Numbering label | i
Input quantity, source of uncertainty | Xi | Uncertainty in spectral power responsivity of reference detector
Estimate of ith input quantity | xi | Values of reference spectral power responsivity
Estimated uncertainty of ith input quantity | ue(xi) | Uncertainty in spectral power responsivity of reference detector
Probability distribution, or spread of the set of values |  | Normal distribution - where values are more likely to fall near the average than further away Rectangular distribution - even spread between the highest and the lowest values U-shaped - where values are more likely to fall near the limits than the centre of the spread
Coverage factor associated with the probability distribution | k | Normal distribution, k = 1 Rectangular distribution, k = sqrt3 U-shaped, k = sqrt2
Standard uncertainty of ith input quantity, estimated uncertainty divided by the coverage factor | u(xi) = ue(xi)/k | Uncertainty in spectral power responsivity of reference detector divided by the covergare factor reported in its calibration certificate
Sensitivity coefficient for ith input quantity, sensitivity of the contribution to the combined uncertainty to a change in the ith input quantity | ci | Degree of change in spectral power responsivity of a detector with a change in wavelength
Contribution to the combined standard uncertainty from the ith input quantity, standard uncertainty multiplied by the sensitivity coefficient | ui(y) = ¦ci¦u(xi)
Combined standard uncertainty, obtained by combining the individual contributions | uc(y)
Expanded uncertainty, product of a combined standard uncertainty and a coverage factor | U | NMISA reports the combined standard uncertainty multiplied by a coverage factor of k = 2.

## UB SR-5416-B
Calibration | No. | Input quantity | Symbol | Estimated uncertainty |  | Probability distribution | Coverage factor | Standard uncertainty |  | Sensitivity coefficient | Uncertainty contribution |  |  | Output quantity | Measurement model | Measurement result
| i | Xi |  | ue(xi) | unit |  | k | u(xi) | unit | ci | ui(y) | unit |  | Y | Y = F(Xi) | Y ± u(y) for a confidence level of 95 %
Absolute | 1 | Absolute spectral power responsivity of the standard detector | uS | 0.44 | % | Normal | 2 | 0.22 | % | 1 | 0.22 | % | 8.1481
Absolute | 2 | Drift in the standard detector | uSd | 0.033531 | % | Normal | 1 | 0.033531 | % | 0.99977 | 0.033523 | % | 1.2416
Absolute | 3 | Spatial uniformity of the standard detector | uSsu | 0.02 | % | Normal | 1 | 0.02 | % | 0.99977 | 0.019995 | % | 0.74057
Absolute | 4 | Linearity of the standard detector | uslin | 0.1 | % | Normal | 2 | 0.05 | % | 0.99977 | 0.049989 | % | 1.8514 | U-shaped | √2
Absolute | 5 | Effect of the amplifier | ua | 0.2 | % | Normal | 2 | 0.1 | % | 0.99977 | 0.099977 | % | 3.7029 | Rectangular | √3
Absolute | 6 | Effect of the multimeter on the detectors' output signal | uAbsM | 0.002 | % | Normal | 2 | 0.001 | % | 0.99977 | 0.00099977 | % | 0.037029
Absolute | 7 | Source stability | uAbsSS | 0.1 | % | Normal | 2 | 0.05 | % | 0.99977 | 0.049989 | % | 1.8514
Absolute | 8 | Stray light and standard detector noise | uSsl | 0.15316 | % | Normal | 1 | 0.15316 | % | 0.99977 | 0.15312 | % | 5.6712
Absolute | 9 | Repeatability of the standard detector | uSr | 0.0051897 | % | Normal | 1 | 0.0051897 | % | 0.99977 | 0.0051886 | % | 0.19217
Absolute | 10 | Spatial uniformity of the test detector | uTsu | 0.1 | % | Normal | 1 | 0.1 | % | 0.99977 | 0.099977 | % | 3.7029
Absolute | 11 | Temperature effects | utemp | 0.01 | % | Normal | 1 | 0.01 | % | 0.99977 | 0.0099977 | % | 0.37029
Absolute | 12 | Repeatability of the test detector | uTr | 0.0069194 | % | Normal | 1 | 0.0069194 | % | 0.99977 | 0.0069178 | % | 0.25622
Spectral | 13 | Pyroelectric detector absorptance | upa | 1.6087 | % | Rectangular | √3 | 0.92877 | % | 0.99977 | 0.92856 | % | 34.391
Spectral | 14 | Effect of the multimeter on the detectors' output signal | uSpecM | 0.002 | % | Normal | 2 | 0.001 | % | 0.99977 | 0.00099977 | % | 0.037029
Spectral | 15 | Source stability | uSpecSS | 0.25 | % | Normal | 2 | 0.125 | % | 0.99977 | 0.12497 | % | 4.6286
Spectral | 16 | Polarisation, diffraction, interference and angular effects | updia | 0.1 | % | Normal | 1 | 0.1 | % | 0.99977 | 0.099977 | % | 3.7029
Spectral | 17 | Spectral wavelength calibration | uλ | 0.25818 | % | Normal | 2 | 0.12909 | % | 0.99977 | 0.12906 | % | 4.7799
Spectral | 18 | Linearity of the test detector | uTlin | 0.1 | % | Rectangular | √3 | 0.057735 | % | 0.99977 | 0.057722 | % | 2.1379
Spectral | 19 | Stray light and test detector noise | uTsl | 0.21787 | % | Normal | 1 | 0.21787 | % | 0.99977 | 0.21782 | % | 8.0674
Spectral | 20 | Repeatability of the test detector | uSpecTr | 0.5533 | % | Normal | 1 | 0.5533 | % | 0.99977 | 0.55318 | % | 20.488
[10 empty] | uc(y) | 1.325 | %
[10 empty] | U (k=2) | 2.7 | %