## Response calculation
λ [Å] |  | Steps | Step interval | Environmental conditions |  | STD - ORW-005.RA | [10 empty] | UUT - PtSi | [5 empty] | STD responsivity [A/W] | STD Uc [%] | UUT responsivity [A/W] | PtSi cert response [A/W] | PtSi cert Uc [A/W] | PtSi cert Uc [%] | Deviation between PtSi cal cert and now [%] | Combined Uc |  | Remarks | BMC for PtSi as STD, Xe source |  |  | Gain: | 1000000
|  |  |  | Temp | Humidity | Light | [6 empty] | Dark |  |  |  | Light |  |  | Dark
[6 empty] | Average [V] |  |  | Std Dev [V] | Std Dev [%] |  |  | Average [V] |  | Std Dev [V] | Net [V] | Average [V] | Std Dev [V] | Std Dev [%] | Average [V] | Std Dev [V] | Net [V] | [10 empty] | 1.81%
Filter no. 1: Blank Xenon arc source |  |  |  | 0.2 | 1.6 | [5 empty] | S/N ratio | [24 empty] | Beam power [W] | Beam power [μW] | Beam power [W] | Beam power [μW]
2000 | 200 | -279093 | 28804.7 | 1 |  | -0.0099425 | 0.0015026 |  | 1.1497e-06 | -0.011563 | 275.7348 | 0.36267 | -0.0095037 | 0.0014362 | 1.1003e-06 | -0.0004388 | -0.013421 | 4.6328e-07 | -0.0034518 | -0.013513 | 6.3059e-07 | 9.1572e-05 | 0.1134 | 6.12 | 0.023665 | 0.02676 | 0.00029 | 1.0837 | 11.5642 | 6.2152 | 0.0014709 | 18/03/2019 | 1.81 | 0.023665 | 11.5642 | 3.8695e-09 | 0.0038695 | 3.8695e-09 | 0.0038695
2100 | 210 | -250288.3 | 28804.7 | 23.6 | 52.2 | -0.01036 | 0.0015656 |  | 2.1869e-06 | -0.02111 | 381.6072 | 0.26205 | -0.0094971 | 0.0014352 | 5.7128e-07 | -0.00086255 | -0.013331 | 5.3452e-07 | -0.0040096 | -0.013514 | 5.458e-07 | 0.00018267 | 0.1163 | 4.25 | 0.024629 | 0.0262 | 0.0002 | 0.76336 | 5.9951 | 4.318 | 0.0010635 |  | 1.65 | 0.024629 | 5.9951 | 7.4166e-09 | 0.0074166 | 7.4166e-09 | 0.0074166
2200 | 220 | -221483.6 | 28804.7 |  |  | -0.012528 | 0.0018933 |  | 1.5399e-05 | -0.12292 | 196.9234 | 0.50781 | -0.0094922 | 0.0014345 | 7.4476e-07 | -0.003036 | -0.012873 | 1.1351e-06 | -0.0088176 | -0.013515 | 5.5469e-07 | 0.0006417 | 0.1209 | 2.29 | 0.025554 | 0.02564 | 0.00017 | 0.66303 | 0.33522 | 2.3841 | 0.00060922 |  |  | 0.025554 | 0.33522 | 2.5112e-08 | 0.025112 | 2.5112e-08 | 0.025112
2300 | 230 | -192678.9 | 28804.7 |  |  | -0.031503 | 0.0047609 |  | 0.00014225 | -0.45154 | 154.7615 | 0.64616 | -0.0094883 | 0.0014339 | 7.6356e-07 | -0.022015 | -0.0089283 | 1.5588e-05 | -0.17459 | -0.013516 | 4.4865e-07 | 0.0045878 | 0.127 | 2.02 | 0.026466 | 0.026 | 0.00016 | 0.61538 | 1.7914 | 2.1117 | 0.00055887 |  |  | 0.026466 | 1.7914 | 1.7335e-07 | 0.17335 | 1.7335e-07 | 0.17335
2400 | 240 | -163874.2 | 28804.7 |  |  | -0.11302 | 0.017079 |  | 0.00073212 | -0.6478 | 141.4131 | 0.70715 | -0.0094843 | 0.0014333 | 5.8971e-07 | -0.10353 | 0.0080205 | 6.1994e-05 | 0.77294 | -0.013517 | 4.5523e-07 | 0.021538 | 0.1304 | 2 | 0.027128 | 0.02677 | 0.00017 | 0.63504 | 1.3355 | 2.0984 | 0.00056924 |  |  | 0.027128 | 1.3355 | 7.9395e-07 | 0.79395 | 7.9395e-07 | 0.79395
2500 | 250 | -135069.5 | 28804.7 |  |  | -0.27913 | 0.042184 |  | 0.0016792 | -0.60157 | 160.5841 | 0.62273 | -0.0094818 | 0.0014329 | 6.5991e-07 | -0.26965 | 0.04219 | 9.2275e-05 | 0.21872 | -0.013519 | 6.246e-07 | 0.055708 | 0.1219 | 1.64 | 0.025184 | 0.02473 | 0.00015 | 0.60655 | 1.8341 | 1.7486 | 0.00044035 |  |  | 0.025184 | 1.8341 | 2.2121e-06 | 2.2121 | 2.2121e-06 | 2.2121
2600 | 260 | -106264.8 | 28804.7 |  |  | -0.5296 | 0.080036 |  | 0.002895 | -0.54664 | 179.6616 | 0.5566 | -0.0094791 | 0.0014325 | 4.3578e-07 | -0.52013 | 0.093717 | 0.00038182 | 0.40742 | -0.013519 | 5.9308e-07 | 0.10724 | 0.1103 | 1.66 | 0.022741 | 0.022396 | 9.8e-05 | 0.43758 | 1.5408 | 1.7167 | 0.0003904 |  |  | 0.022741 | 1.5408 | 4.7155e-06 | 4.7155 | 4.7155e-06 | 4.7155
2700 | 270 | -77460.1 | 28804.7 |  |  | -0.86818 | 0.1312 |  | 0.0046691 | -0.53781 | 183.9103 | 0.54374 | -0.0094776 | 0.0014323 | 6.1721e-07 | -0.8587 | 0.16354 | 0.000586 | 0.35832 | -0.013521 | 5.5695e-07 | 0.17706 | 0.1001 | 1.64 | 0.02064 | 0.02025 | 0.0001 | 0.49383 | 1.9283 | 1.7127 | 0.00035352 |  |  | 0.02064 | 1.9283 | 8.5784e-06 | 8.5784 | 8.5784e-06 | 8.5784
2800 | 280 | -48655.4 | 28804.7 |  |  | -1.2794 | 0.19335 |  | 0.013756 | -1.0752 | 92.319 | 1.0832 | -0.0094757 | 0.001432 | 5.5096e-07 | -1.2699 | 0.25876 | 0.0014241 | 0.55036 | -0.01352 | 4.6737e-07 | 0.27228 | 0.1025 | 1.63 | 0.021977 | 0.02146 | 0.00013 | 0.60578 | 2.4086 | 1.7389 | 0.00038216 |  |  | 0.021977 | 2.4086 | 1.2389e-05 | 12.3894 | 1.2389e-05 | 12.3894
2900 | 290 | -19850.7 | 28804.7 |  |  | -1.862 | 0.28139 |  | 0.018634 | -1.0007 | 99.4169 | 1.0059 | -0.0094746 | 0.0014318 | 3.5741e-07 | -1.8525 | 0.41808 | 0.001489 | 0.35615 | -0.01352 | 4.2849e-07 | 0.4316 | 0.1147 | 1.64 | 0.026723 | 0.02618 | 0.00013 | 0.49656 | 2.0754 | 1.7135 | 0.00045791 |  |  | 0.026723 | 2.0754 | 1.6151e-05 | 16.1508 | 1.6151e-05 | 16.1508
3000 | 300 | 8954 | 28804.7 |  |  | -2.5509 | 0.3855 |  | 0.02538 | -0.99495 | 100.1342 | 0.99866 | -0.0094731 | 0.0014316 | 6.0211e-07 | -2.5414 | 0.59766 | 0.0019461 | 0.32561 | -0.013521 | 4.2444e-07 | 0.61118 | 0.1295 | 1.63 | 0.031144 | 0.03063 | 0.00011 | 0.35913 | 1.6772 | 1.6691 | 0.00051982 |  |  | 0.031138 | 1.6583 | 1.9625e-05 | 19.6246 | 1.9625e-05 | 19.6246
3100 | 310 | 37758.7 | 28804.7 |  |  | -3.139 | 0.47438 |  | 0.03824 | -1.2182 | 81.8391 | 1.2219 | -0.0094712 | 0.0014313 | 5.3329e-07 | -3.1295 | 0.74418 | 0.003172 | 0.42625 | -0.01352 | 2.0741e-07 | 0.7577 | 0.1375 | 1.63 | 0.03329 | 0.032732 | 9.5e-05 | 0.29024 | 1.7056 | 1.6556 | 0.00055117 |  |  | 0.033286 | 1.6934 | 2.276e-05 | 22.7603 | 2.276e-05 | 22.7603
3200 | 320 | 66563.4 | 28804.7 |  |  | -3.6094 | 0.54547 |  | 0.036751 | -1.0182 | 97.9563 | 1.0209 | -0.0094699 | 0.0014311 | 4.8885e-07 | -3.6 | 0.85132 | 0.0029152 | 0.34243 | -0.01352 | 3.2257e-07 | 0.86484 | 0.1422 | 1.63 | 0.034161 | 0.033631 | 9.6e-05 | 0.28545 | 1.5772 | 1.6548 | 0.00056531 |  |  | 0.03416 | 1.5727 | 2.5316e-05 | 25.3161 | 2.5316e-05 | 25.3161
3300 | 330 | 95368.1 | 28804.7 |  |  | -4.0329 | 0.60946 |  | 0.041553 | -1.0304 | 96.8245 | 1.0328 | -0.009469 | 0.001431 | 7.4787e-07 | -4.0234 | 0.95236 | 0.0035652 | 0.37435 | -0.013521 | 5.4884e-07 | 0.96588 | 0.1455 | 1.63 | 0.03493 | 0.034394 | 8.9e-05 | 0.25877 | 1.5572 | 1.6504 | 0.00057648 |  |  | 0.034931 | 1.5614 | 2.7652e-05 | 27.6522 | 2.7652e-05 | 27.6522
3400 | 340 | 124172.8 | 28804.7 |  |  | -4.4235 | 0.66849 |  | 0.044489 | -1.0057 | 99.2158 | 1.0079 | -0.0094674 | 0.0014308 | 7.2127e-07 | -4.414 | 1.0553 | 0.0060864 | 0.57675 | -0.013522 | 6.7262e-07 | 1.0688 | 0.1478 | 1.63 | 0.035789 | 0.03534 | 9e-05 | 0.25467 | 1.2702 | 1.6498 | 0.00059044 |  |  | 0.035793 | 1.281 | 2.9865e-05 | 29.8646 | 2.9865e-05 | 29.8646
3500 | 350 | 152977.5 | 28804.7 |  |  | -4.6946 | 0.70947 |  | 0.048847 | -1.0405 | 95.9137 | 1.0426 | -0.0094675 | 0.0014308 | 5.498e-07 | -4.6851 | 1.1374 | 0.0059386 | 0.52214 | -0.013522 | 5.7128e-07 | 1.1509 | 0.1481 | 1.63 | 0.03638 | 0.035982 | 9.2e-05 | 0.25568 | 1.1061 | 1.6499 | 0.00060024 |  |  | 0.036384 | 1.1179 | 3.1635e-05 | 31.635 | 3.1635e-05 | 31.635
3600 | 360 | 181782.2 | 28804.7 |  |  | -4.7776 | 0.72201 |  | 0.049975 | -1.046 | 95.4099 | 1.0481 | -0.009467 | 0.0014307 | 3.9679e-07 | -4.7681 | 1.1872 | 0.0058361 | 0.49158 | -0.013523 | 6.5061e-07 | 1.2007 | 0.1447 | 1.63 | 0.036439 | 0.03585 | 0.0001 | 0.27894 | 1.6441 | 1.6537 | 0.0006026 |  |  | 0.036442 | 1.6514 | 3.2952e-05 | 32.9518 | 3.2952e-05 | 32.9518
3700 | 370 | 210586.9 | 28804.7 |  |  | -4.9434 | 0.74707 |  | 0.053122 | -1.0746 | 92.8796 | 1.0767 | -0.009466 | 0.0014305 | 4.5979e-07 | -4.9339 | 1.4272 | 0.0090548 | 0.63443 | -0.013523 | 5.7304e-07 | 1.4407 | 0.1449 | 1.33 | 0.042312 | 0.04117 | 0.00026 | 0.63153 | 2.7739 | 1.4723 | 0.00062297 |  |  | 0.042311 | 2.7716 | 3.4051e-05 | 34.0505 | 3.4051e-05 | 34.0505
3800 | 380 | 239391.6 | 28804.7 |  |  | -5.4795 | 0.82808 |  | 0.057237 | -1.0446 | 95.5676 | 1.0464 | -0.0094657 | 0.0014305 | 6.3601e-07 | -5.47 | 2.1375 | 0.014545 | 0.68048 | -0.013523 | 3.9225e-07 | 2.151 | 0.1567 | 0.79 | 0.061621 | 0.0604 | 0.00042 | 0.69536 | 2.021 | 1.0524 | 0.00064852 |  |  | 0.061612 | 2.0064 | 3.4908e-05 | 34.9076 | 3.4908e-05 | 34.9076
3900 | 390 | 268196.3 | 28804.7 |  |  | -6.0409 | 0.91292 |  | 0.068257 | -1.1299 | 88.3631 | 1.1317 | -0.0094648 | 0.0014304 | 5.8565e-07 | -6.0314 | 2.8938 | 0.01298 | 0.44855 | -0.013523 | 7.0415e-07 | 2.9073 | 0.1691 | 0.65 | 0.081511 | 0.08051 | 0.00036 | 0.44715 | 1.2434 | 0.78895 | 0.00064308 |  |  | 0.081489 | 1.2163 | 3.5668e-05 | 35.6676 | 3.5668e-05 | 35.6676
4000 | 400 | 297001 | #REF! |  |  | -6.6171 | 1 |  | 0.070879 | -1.0712 | 93.2238 | 1.0727 | -0.0094643 | 0.0014303 | 3.3147e-07 | -6.6076 | 3.5475 | 0.017099 | 0.48202 | -0.013523 | 5.7967e-07 | 3.561 | 0.1798 | 0.49 | 0.096899 | 0.09599 | 0.00032 | 0.33337 | 0.94651 | 0.59265 | 0.00057427 |  |  | 0.096864 | 0.91017 | 3.675e-05 | 36.7497 | 3.675e-05 | 36.7497
λ [Å] |  | Steps | Step interval | Environmental conditions |  | STD - ORW-005.RA | [10 empty] | UUT - PtSi | [5 empty] | STD responsivity [A/W] | STD Uc [%] | UUT responsivity [A/W] | PtSi cert response [A/W] | PtSi cert Uc [A/W] | PtSi cert Uc [%] | Deviation between PtSi cal cert and now [%] | Combined Uc |  | Remarks
|  |  |  | Temp | Humidity | Light | [6 empty] | Dark |  |  |  | Light |  |  | Dark
[6 empty] | Average [V] |  |  | Std Dev [V] | Std Dev [%] |  |  | Average [V] |  | Std Dev [V] | Net [V] | Average [V] | Std Dev [V] | Std Dev [%] | Average [V] | Std Dev [V] | Net [V]
Filter no. 1: Blank LDLS |  |  |  | 0.2 | 1.6 |  |  | % diff in signals |  |  | S/N ratio
2000 | 200 | -279093 | 28804.7 | 1 |  | 0.23136 | 0.032595 | 3.1092 | 0.00050567 | 0.21856 | 454.885 | 0.21984 | 0.0094873 | 0.0013366 | 1.4687e-06 | 0.22187 | -0.039875 | 0.00017453 | -0.43768 | 0.013413 | 6.9582e-07 | -0.053289 | 0.1134 | 6.12 | 0.027236 | 0.02676 | 0.00029 | 1.0837 | 1.78 | 6.2152 | 0.0016928 | 18/03/2019 |  |  |  | 1.9565e-06 | 1.9565
2100 | 210 | -250288.3 | 28804.7 | 23.5 | 54.8 | 0.66439 | 0.093603 | 9.2037 | 0.0041148 | 0.61933 | 161.1413 | 0.62057 | 0.0094771 | 0.0013352 | 1.8406e-06 | 0.65492 | -0.1343 | 0.0011114 | -0.82755 | 0.013415 | 1.6025e-06 | -0.14771 | 0.1163 | 4.25 | 0.026231 | 0.0262 | 0.0002 | 0.76336 | 0.11796 | 4.318 | 0.0011327 |  |  |  |  | 5.6313e-06 | 5.6313
2200 | 220 | -221483.6 | 28804.7 |  |  | 1.5263 | 0.21504 | 21.3146 | 0.016643 | 1.0904 | 91.6322 | 1.0913 | 0.0094703 | 0.0013342 | 2.6754e-06 | 1.5169 | -0.31426 | 0.002319 | -0.73791 | 0.013415 | 1.5922e-06 | -0.32768 | 0.1209 | 2.29 | 0.026117 | 0.02564 | 0.00017 | 0.66303 | 1.8598 | 2.3841 | 0.00062264 |  |  |  |  | 1.2547e-05 | 12.5466
2300 | 230 | -192678.9 | 28804.7 |  |  | 3.1198 | 0.43954 | 43.4778 | 0.033958 | 1.0885 | 91.8333 | 1.0889 | 0.0094653 | 0.0013335 | 1.9459e-06 | 3.1104 | -0.64288 | 0.0045682 | -0.71059 | 0.013416 | 1.8241e-06 | -0.6563 | 0.127 | 2.02 | 0.026797 | 0.026 | 0.00016 | 0.61538 | 3.0669 | 2.1117 | 0.00056587 |  |  |  |  | 2.4491e-05 | 24.4912
2400 | 240 | -163874.2 | 28804.7 |  |  | 4.1405 | 0.58334 | 56.6259 | 0.046567 | 1.1247 | 88.8872 | 1.125 | 0.0094612 | 0.0013329 | 1.2281e-06 | 4.1311 | -0.85327 | 0.006007 | -0.704 | 0.013416 | 1.4171e-06 | -0.86669 | 0.1304 | 2 | 0.027357 | 0.02677 | 0.00017 | 0.63504 | 2.1945 | 2.0984 | 0.00057407 |  |  |  |  | 3.168e-05 | 31.68
2500 | 250 | -135069.5 | 28804.7 |  |  | 4.8551 | 0.68401 | 64.1822 | 0.057181 | 1.1778 | 84.8833 | 1.1781 | 0.0094581 | 0.0013325 | 2.2417e-06 | 4.8456 | -1.0012 | 0.0071798 | -0.71712 | 0.013415 | 1.0477e-06 | -1.0146 | 0.1219 | 1.64 | 0.025524 | 0.02473 | 0.00015 | 0.60655 | 3.2123 | 1.7486 | 0.00044631 |  |  |  |  | 3.9751e-05 | 39.7508
2600 | 260 | -106264.8 | 28804.7 |  |  | 5.3006 | 0.74678 | 66.6742 | 0.068721 | 1.2965 | 77.1131 | 1.2968 | 0.0094557 | 0.0013322 | 1.9234e-06 | 5.2912 | -1.0869 | 0.012722 | -1.1704 | 0.013415 | 2.878e-06 | -1.1004 | 0.1103 | 1.66 | 0.022938 | 0.022396 | 9.8e-05 | 0.43758 | 2.4202 | 1.7167 | 0.00039378 |  |  |  |  | 4.7971e-05 | 47.9708
2700 | 270 | -77460.1 | 28804.7 |  |  | 5.9585 | 0.83946 | 70.8254 | 0.070069 | 1.176 | 85.0182 | 1.1762 | 0.0094537 | 0.0013319 | 2.6459e-06 | 5.949 | -1.218 | 0.013737 | -1.1278 | 0.013415 | 1.6119e-06 | -1.2315 | 0.1001 | 1.64 | 0.020721 | 0.02025 | 0.0001 | 0.49383 | 2.3258 | 1.7127 | 0.0003549 |  |  |  |  | 5.9431e-05 | 59.4307
2800 | 280 | -48655.4 | 28804.7 |  |  | 7.098 | 1 | 80.6653 | 0.0824 | 1.1609 | 86.1242 | 1.1611 | 0.0094535 | 0.0013319 | 1.4821e-06 | 7.0885 | -1.5389 | 0.018072 | -1.1744 | 0.013413 | 3.0679e-06 | -1.5523 | 0.1025 | 1.63 | 0.022446 | 0.02146 | 0.00013 | 0.60578 | 4.5942 | 1.7389 | 0.00039032 |  |  |  |  | 6.9157e-05 | 69.1566
2900 | 290 | -19850.7 | 28804.7 |  |  | 0.097344 | 0.013714 | -26.7675 | 0.00017896 | 0.18385 | 536.4975 | 0.18639 | 0.0094422 | 0.0013303 | 7.4449e-07 | 0.087902 | -0.011313 | 8.1043e-05 | -0.71635 | 0.0094772 | 8.3535e-07 | -0.02079 | 0.1147 | 1.64 | 0.027129 | 0.02618 | 0.00013 | 0.49656 | 3.6239 | 1.7135 | 0.00046486 |  |  |  |  | 7.6636e-07 | 0.76636
3000 | 300 | 8954 | 28804.7 |  |  | 0.11243 | 0.01584 | -36.9657 | 0.00080854 | 0.71914 | 137.4106 | 0.72775 | 0.0094414 | 0.0013302 | 5.033e-07 | 0.10299 | -0.015589 | 0.00010012 | -0.64223 | 0.0094767 | 6.6196e-07 | -0.025066 | 0.1295 | 1.63 | 0.031517 | 0.03063 | 0.00011 | 0.35913 | 2.8972 | 1.6691 | 0.00052605 |  |  |  |  | 7.9529e-07 | 0.79529
3100 | 310 | 37758.7 | 28804.7 |  |  | 0.11925 | 0.0168 | -45.7581 | 0.00085798 | 0.71949 | 137.437 | 0.72761 | 0.0094417 | 0.0013302 | 9.9524e-07 | 0.10981 | -0.017303 | 0.00010012 | -0.57864 | 0.0094769 | 2.0363e-06 | -0.02678 | 0.1375 | 1.63 | 0.033534 | 0.032732 | 9.5e-05 | 0.29024 | 2.449 | 1.6556 | 0.0005552 |  |  |  |  | 7.986e-07 | 0.7986
3200 | 320 | 66563.4 | 28804.7 |  |  | 0.12185 | 0.017167 | -52.8306 | 0.00089382 | 0.73354 | 134.8379 | 0.74163 | 0.0094415 | 0.0013302 | 6.4526e-07 | 0.11241 | -0.017761 | 0.00013606 | -0.76606 | 0.0094766 | 2.2458e-06 | -0.027238 | 0.1422 | 1.63 | 0.034456 | 0.033631 | 9.6e-05 | 0.28545 | 2.4533 | 1.6548 | 0.00057018 |  |  |  |  | 7.905e-07 | 0.7905
3300 | 330 | 95368.1 | 28804.7 |  |  | 0.12449 | 0.017538 | -59.1926 | 0.0009024 | 0.7249 | 136.4757 | 0.73273 | 0.0094406 | 0.00133 | 8.904e-07 | 0.11505 | -0.018208 | 0.00013094 | -0.71912 | 0.0094772 | 2.1083e-06 | -0.027685 | 0.1455 | 1.63 | 0.035014 | 0.034394 | 8.9e-05 | 0.25877 | 1.8018 | 1.6504 | 0.00057787 |  |  |  |  | 7.9069e-07 | 0.79069
3400 | 340 | 124172.8 | 28804.7 |  |  | 0.12533 | 0.017657 | -65.0835 | 0.00093844 | 0.74879 | 132.1311 | 0.75682 | 0.0094399 | 0.0013299 | 1.6009e-06 | 0.11589 | -0.018776 | 0.00013858 | -0.73805 | 0.0094768 | 1.2859e-06 | -0.028253 | 0.1478 | 1.63 | 0.036033 | 0.03534 | 9e-05 | 0.25467 | 1.9619 | 1.6498 | 0.00059447 |  |  |  |  | 7.8408e-07 | 0.78408
3500 | 350 | 152977.5 | 28804.7 |  |  | 0.12303 | 0.017333 | -69.2136 | 0.00090974 | 0.73943 | 133.7768 | 0.74751 | 0.0094392 | 0.0013298 | 2.7644e-06 | 0.11359 | -0.018519 | 0.00010928 | -0.5901 | 0.0094771 | 6.1238e-07 | -0.027996 | 0.1481 | 1.63 | 0.0365 | 0.035982 | 9.2e-05 | 0.25568 | 1.4393 | 1.6499 | 0.00060222 |  |  |  |  | 7.6701e-07 | 0.76701
3600 | 360 | 181782.2 | 28804.7 |  |  | 0.11723 | 0.016516 | -70.5495 | 0.00086723 | 0.73974 | 133.6485 | 0.74823 | 0.0094402 | 0.00133 | 1.0148e-06 | 0.10779 | -0.017911 | 0.00011235 | -0.62725 | 0.0094763 | 1.3809e-06 | -0.027388 | 0.1447 | 1.63 | 0.036765 | 0.03585 | 0.0001 | 0.27894 | 2.5516 | 1.6537 | 0.00060798 |  |  |  |  | 7.4495e-07 | 0.74495
3700 | 370 | 210586.9 | 28804.7 |  |  | 0.11493 | 0.016192 | -73.0875 | 0.00088684 | 0.77163 | 128.0958 | 0.78067 | 0.0094409 | 0.0013301 | 8.66e-07 | 0.10549 | -0.022677 | 0.00015259 | -0.6729 | 0.0094773 | 2.9885e-06 | -0.032154 | 0.1449 | 1.33 | 0.044166 | 0.04117 | 0.00026 | 0.63153 | 7.2778 | 1.4723 | 0.00065027 |  |  |  |  | 7.2802e-07 | 0.72802
3800 | 380 | 239391.6 | 28804.7 |  |  | 0.12112 | 0.017064 | -81.1019 | 0.00092592 | 0.76446 | 129.3749 | 0.77295 | 0.0094411 | 0.0013301 | 1.0573e-06 | 0.11168 | -0.036418 | 0.0002536 | -0.69637 | 0.0094775 | 2.9177e-06 | -0.045895 | 0.1567 | 0.79 | 0.064396 | 0.0604 | 0.00042 | 0.69536 | 6.6159 | 1.0524 | 0.00067773 |  |  |  |  | 7.127e-07 | 0.7127
3900 | 390 | 268196.3 | 28804.7 |  |  | 0.12618 | 0.017777 | -89.5143 | 0.00097661 | 0.77397 | 127.8427 | 0.78221 | 0.0094412 | 0.0013301 | 8.093e-07 | 0.11674 | -0.047999 | 0.00033087 | -0.68932 | 0.0094785 | 2.7066e-06 | -0.057478 | 0.1691 | 0.65 | 0.083256 | 0.08051 | 0.00036 | 0.44715 | 3.4111 | 0.78895 | 0.00065685 |  |  |  |  | 6.9037e-07 | 0.69037
4000 | 400 | 297001 | #REF! |  |  | 0.12705 | 0.0179 | -98.21 | 0.0009337 | 0.7349 | 134.6476 | 0.74268 | 0.0094412 | 0.0013301 | 8.4998e-07 | 0.11761 | -0.054427 | 0.00033536 | -0.61616 | 0.0094778 | 2.7295e-06 | -0.063904 | 0.1798 | 0.49 | 0.097696 | 0.09599 | 0.00032 | 0.33337 | 1.7773 | 0.59265 | 0.000579 |  |  |  |  | 6.5412e-07 | 0.65412
[13 empty] | y = 288.047x-855187 |  |  |  | Step calculator | Wavelength [A] | Step count
[13 empty] | 288.047 |  |  | 968150.51 |  | 9050 | 1751638.35
855187
[9 empty] | 2000 | -279093 |  |  | 5500 |  | 729071.5 | 288047
[5 empty] | 6970 |  |  |  | 2500 | -135069.5 |  |  | 5000 |  | 585048
[13 empty] | 4500 |  | 441024.5 | 288047
[13 empty] | 3500 |  | 152977.5
[5 empty] | 9658 |  |  |  | -144023.5 |  |  |  | 9618 |  | 1915249.046
[13 empty] | 9698 |  | 1938292.806
[10 empty] | Si | PtSi
[9 empty] | 2000 | 0.0004388 | 9.1572e-05
[9 empty] | 2100 | 0.00086255 | 0.00018267
[9 empty] | 2200 | 0.003036 | 0.0006417
[9 empty] | 2300 | 0.022015 | 0.0045878
[9 empty] | 2400 | 0.10353 | 0.021538
[9 empty] | 2500 | 0.26965 | 0.055708
[9 empty] | 2600 | 0.52013 | 0.10724
[9 empty] | 2700 | 0.8587 | 0.17706
[9 empty] | 2800 | 1.2699 | 0.27228
[9 empty] | 2900 | 1.8525 | 0.4316
[9 empty] | 3000 | 2.5414 | 0.61118
[9 empty] | 3100 | 3.1295 | 0.7577
[9 empty] | 3200 | 3.6 | 0.86484
[9 empty] | 3300 | 4.0234 | 0.96588
[9 empty] | 3400 | 4.414 | 1.0688
[9 empty] | 3500 | 4.6851 | 1.1509
[9 empty] | 3600 | 4.7681 | 1.2007
[9 empty] | 3700 | 4.9339 | 1.4407
[9 empty] | 3800 | 5.47 | 2.151
[9 empty] | 3900 | 6.0314 | 2.9073
[9 empty] | 4000 | 6.6076 | 3.561

## Uncertainty calc
Uncertainty contributions | [18 empty] | Calibration of spectral power responsivity of STD detector | [7 empty] | Drift in STD, from paper | Drift in STD | Polarization, non-uniform beam profile and repeatability | Polarization, non-uniform beam profile and repeatability | Beam divergence & vignetting divergence, negligible if 2,4° | Linearity of STD | Spectral wavelength calibration | Lamp stability / Drift during cal. |  |  |  |  | Temperature dependence of STD | DVM uncertainty |  |  |  |  | Spatial uniformity of UUT | Amplifier uncertainty same gain setting used for STD & UUT | Scattered light (internal stray light) | External stray light (dark reading subtracted, use as indication of signal-noise ratio) | Repeatability in STD |  |  |  | Resolution of STD |  | Bandwidth effects on UUT | Spatial uniformity of UUT |  |  |  | Temperature dependence of UUT |  | DVM uncertainty |  |  |  |  | Repeatability in UUT |  |  |  | Resolution of UUT
Remarks | [30 empty] | CCPR-K2.c 2003 | [8 empty] | Uncertainty |  | Accuracy: ± (0,002 % of reading + 0,0006 % of range) | [19 empty] | Uncertainty |  | Accuracy: ± (0,002 % of reading + 0,0006 % of range)
Estimated U | [18 empty] | 0,49 to 6,12 % | [9 empty] | 0.041% | 0.01% |  |  | 0,3 nm | 0.217% |  |  |  |  | See below | Range [V] | Uc [V] | Range [V] | % of reading | % of range |  | 0% | 0.0015% |  | See below |  |  |  | See below | [6 empty] | See below |  | Range [V] | Uc [V] | Range [V] | % of reading | % of range | See below |  |  |  | See below
Probability distr. | [18 empty] | Normal | [7 empty] | Normal | Normal | Normal | Normal | Normal | Normal | Normal | Rectangular |  |  |  |  | Normal | 0.0001 | 1.2e-06 | 0.1 | 0.003% | 0.003% | Normal | Normal | normal | normal | Normal |  |  |  | Rectangular |  | Rectangular | Normal |  |  |  | Rectangular |  | 0.0001 | 1.2e-06 | 0.1 | 0.003% | 0.003% | Normal |  |  |  | Rectangular
k = | [18 empty] | 2 | [7 empty] | 3.18 | 1 | 1 | 1 | 1 | 1 | 2 | 1.7321 |  |  |  |  | 1 | 0.001 | 1.2e-06 | 1 | 0.002% | 0.0006% | 1 | 1 | 1 | 1 | 1 |  |  |  | 1.7321 |  | 1.7321 | 1 |  |  |  | 1.7321 |  | 0.001 | 1.2e-06 | 1 | 0.002% | 0.0006% | 1 |  |  |  | 1.7321
Standard U | [18 empty] | See below | [7 empty] | See below | See below | See below | See below | See below | See below | See below | See below | [5 empty] | 0.01 | 1.2e-06 | 10 | 0.0015% | 0.0004% | See below | See below | See below | See below | See below |  |  |  | See below |  | See below | See below | [5 empty] | 0.01 | 1.2e-06 | 10 | 0.0015% | 0.0004% | See below |  |  |  | See below
Max std Uc [%] | [18 empty] | 0.00072416 | [14 empty] | 0.21696 | [5 empty] | 0.1 | 1.2e-06 | divisor |  | 1.7321 | 0.086 | 0 | 0.153% | 95.5867 | 1.2182 |  |  |  | 2.7865e-07 | [8 empty] | 0.1 | 1.2e-06 | divisor |  | 1.7321 | 0.77294 |  |  |  | 2.7355e-06
[40 empty] | 1 | 6e-06 |  |  | See below | [17 empty] | 1 | 6e-06 |  |  | See below
[40 empty] | 10 | 4e-05 |  |  | 0.13674 | [17 empty] | 10 | 4e-05 |  |  | 0.65522
[40 empty] | divisor | 2 | [20 empty] | divisor | 2
| Standard - Si: ORW-005.RA | [7 empty] | UUT - PtSi: ORW-523.RA | [8 empty] | Standard and Reference Equipment (Uncorrelated) | [35 empty] | UUT Uncertainty contributions | [17 empty] | Uncertainty and remarks
| Spectral power responsivity | Temp coeff interpolated | Corrected Spectral power responsivity | Deviation | Signal of STD (Light minus dark) | Standard Deviation in signal (Light) | S/N |  | Spectral power responsivity | PtSi spectral power responsivity | Spectral power responsivity calibration values from PTB of PtSi, corrected for deviation from calibration temperature (25,2 °C), using temperature coefficient. ΔT = 25,2-23,5 |  | Signal of UUT (Light minus dark) | Standard Deviation in signal (Light) | S/N |  |  | Absolute spectral power responsivity of the standard detector | [6 empty] | Drift in STD |  |  |  | Polarization, non-uniform beam profile and repeatability | Inter-reflections between detector and output aperture of the monochromator | Beam divergence & vignetting divergence < 2,4° | Linearity of STD | Spectral wavelength calibration | Lamp stability / Drift during cal. |  | Temperature coefficient of STD |  | Temperature dependence of STD |  | DVM uncertainty - STD |  |  |  |  | Spatial uniformity of STD | Amplifier uncertainty | Scattered light (internal stray light) | External stray light - STD | Repeatability uncertainty contribution | Resolution of STD |  |  |  | Bandwidth effects on UUT |  | Spatial uniformity of UUT | Temperature coefficient of UUT |  | Temperature dependence of UUT |  | Linearity of UUT | DVM uncertainty - UUT |  |  |  |  | Repeatability uncertainty contribution | Resolution of UUT |  |  |  | Combined uncertainty standard contribution BMC (k = 1) | Relative Expanded Uncertainty BMC (k = 2) | Combined uncertainty contribution UUT included (k = 1) | Expanded Uncertainty UUT included (k = 2) |  | Deviation between PTB calibration & NMISA calculated spectral power responsivity | Verification statement | Remarks |  |  |  | Deviation between temp uncorr PTB calibration & NMISA calculated spectral power responsivity
[8 empty] | CIE 202:2011 page 7 |  |  |  | Deviation [%] |  |  |  | % noise of signal, insignificant if less than 1/3 of uncertainty | CIE 202:2011 page 7 | Spectral power responsivity Certificate Uncertainty (k=2) | Spectral power responsivity Certificate Uncertainty (k=1) | Spectral power responsivity Certificate Uncertainty | Spectral power responsivity certificate standard uncertainty | Sensitivity Coefficient | Standard uncertainty contribution | Standard uncertainty contribution | Drift in response (per annum) | Standard deviation of drift rate σ(λ) (Prediction contribution U p.a.) | for k = 3,18 | Drift including drift rate, worst case. k = 1 | Standard uncertainty contribution | Standard uncertainty contribution | Standard uncertainty contribution | Standard uncertainty contribution | Standard uncertainty contribution | Standard uncertainty contribution | ? | Standard uncertainty contribution |  | Standard uncertainty contribution |  | Standard uncertainty |  | Standard uncertainty contribution for accuracy |  | Combined standard uncertainty contribution for DVM | Standard uncertainty contribution | Standard uncertainty contribution | Internal stray light uncertainty estimate | External stray light uncertainty estimate |  | Resolution of STD | Semi range of resolution of STD | Resolution of STD standard uncertainty contribution | Resolution of STD standard uncertainty contribution | Standard uncertainty contribution |  | Standard uncertainty contribution | Standard uncertainty contribution |  | Standard uncertainty contribution |  | Standard uncertainty contribution | Standard uncertainty |  | Standard uncertainty contribution for accuracy |  | Combined standard uncertainty contribution for DVM |  | Resolution of UUT | Semi range of resolution of UUT | Resolution of UUT standard uncertainty contribution | Resolution of STD standard uncertainty contribution
Source/ reference | NPL Certificate 2017060356 | Interpolated from CCPR-k2.b 2004 report | ΔT = 2 °C |  | Measured | Calculated | Kostkowski | 100 S/N ratio = 1 % Uc | Calculated | PTB Certificate 73328 18 PTB |  |  | Measured | Calculated | Kostkowski | Kostkowski page 276 | 100 S/N ratio = 1 % Uc | NPL Certificate 2017060356 | [6 empty] | From paper: Predicting wavlength dependent responsivity-drift is silicon, Littler et al (NewRad proceedings) |  |  |  | CCPR-K2.c 2003 worst case from participants Uaver p50 | CCPR-K2.c 2003 | PTB Uncertainty values | CCPR-K2.c 2003 | Certificate OR\SR-5491 & calculated | Measured |  | CCPR-K2.c 2003, Si detector manufacturer specifications & estimate |  | Optical Radiometry book & estimate |  | k=2 | k=1 | k=1,732 | k=1 | DVM certificate and accuracy specification | CCPR-K2.c 2003 | Common factor for both STD & UUT, therefore no contribution | McPherson specification | Calculated from dark reading of std | Caclulated from measured data | Caclulated from measured data |  |  |  | CCPR-K2.c, detector of same type |  | CCPR-K2.c 2003 | CCPR-K2.c 2003 |  | CCPR-K2.c 2003 |  | CCPR-K2.c 2003 | k=2 | k=1 | k=1,732 | k=1 | DVM certificate and accuracy specification | Caclulated from measured data | Caclulated from measured data | [7 empty] | rounded up
Probability ditribution | [17 empty] | Gaussian
Coverage factor, k | [17 empty] | k = 2 | k = 1
| sSTD_ [A/W] |  |  |  | VSTD_ [V] | [V] |  | rel Uc | sUUT_ [A/W] | [A/W] | [A/W] |  | VUUT_ [V] | [V] |  |  | rel Uc | Us [%] | [6 empty] | [%] | [%] | [%] | U(Drift) [%] | U(Pol) [%] | U(Refl) [%] | U(BD&V) [%] | U(Lin) [%] | Uc(λ) [%] | Uc(S) [%] |  | ΔT [°C] | -2.5 | U(Td) [%] | U(Td) [%] | U(DVM_u) [%] | U(DVM_u) [%] | U(DVM_acc) [%] | U(DVM_acc) [%] | U(DVM) [%] | U(SU) [%] | u(Amp) [%] | U(SLi) [%] | U(SLe) [%] | USTD(esdm) [%] | [V] | [V] | [V] | U(ResSTD) [%] | U(bw) [%] | U(bw) [%] | U(SU) [%] | [%/deg C] | ΔT = 25,2-23,5 [°C] | U(T) [%] | U(T) [%] | U(Lin) [%] | U(DVM_u) [%] | U(DVM_u) [%] | U(DVM_acc) [%] | U(DVM_acc) [%] | U(DVM) [%] | [%] | [V] | [V] | [V] | U(ResUUT) [%] | U [%] | U [%] | U [%] | U [%] |  | [%] | [5 empty] | [%]
Wavelength [nm] |  | [1/K] | [A/W] | [%] |  |  | [1] | [%] | [6 empty] | [1] |  | [%] |  |  | U(sSTD) [A/W] | U(sSTD) [A/W] | c_sUUT,sSTD [1] | Uc(sSTD) [A/W] | Uc(sSTD) [%] | [11 empty] | [%/deg C] | U(Tc) [%] | [20 empty] | 1.7 | [21 empty] | Resolution |  | En
200 | 0.1134 |  |  |  | 0.0004388 | 1.1497e-06 | 275.7348 | 0.36267 | 0.023665 | 0.02676 | 0.026719 | 0.153 | 9.1572e-05 | 4.6328e-07 | 117.0282 | 0.85449 | 0.85449 | 6.12 | 3.06 | 0.0069401 | 0.00347 | 0.20869 | 0.00072416 | 0.63859 | -0.01076 | 0.05206 | 0.16555 | 0.05316 | 0.041 | 0.01 | 0.005 | 0.02 | 0.061189 | 0.21696 | 494.4341 | 0.05 | -0.125 | 0.02 | 0.02 | 0.27347 | 0.13674 | 3.0132e-06 | 1.7397e-06 | 0.13674 | 0.086 | 0 | 0.00153 | 95.5867 | 0.011563 | 1e-11 | 5e-12 | 2.8868e-12 | 2.9034e-08 | 0.029 | 0.016743 | 0.124% | -0.09 | -0.153 | 0.016% | 0.009238% | 0.02 | 1.3104 | 0.65522 | 3.0027e-06 | 1.7336e-06 | 0.65522 | -0.0034518 | 1e-10 | 5e-11 | 2.8868e-11 | -2.1508e-07 | 95.6361 | 191.2723 | 95.6385 | 191.277 | 191.3 | 11.4287 | OK | This range is in noise level, will not use these values | 0.045272 | 191.3 | 0.06745 | 11.5642
210 | 0.1163 |  |  |  | 0.00086255 | 2.1869e-06 | 381.6072 | 0.26205 | 0.024629 | 0.0262 | 0.026155 | 0.17 | 0.00018267 | 5.3452e-07 | 239.1074 | 0.41822 | 0.41822 | 4.25 | 2.125 | 0.0049428 | 0.0024714 | 0.21177 | 0.00052337 | 0.45002 | -0.012099 | 0.050125 | 0.1594 | 0.051564 | 0.041 | 0.01 | 0.005 | 0.02 | 0.058521 | 0.21696 | 251.5288 | 0.05 | -0.125 | 0.02 | 0.02 | 0.13912 | 0.069561 | 3.0259e-06 | 1.747e-06 | 0.069561 | 0.086 | 0 | 0.00153 | 91.674 | 0.02111 | 1e-10 | 5e-11 | 2.8868e-11 | 2.7865e-07 | 0.015 | 0.0086603 | 0.124% | -0.1 | -0.17 | 0.017% | 0.009815% | 0.02 | 0.65694 | 0.32847 | 3.0055e-06 | 1.7352e-06 | 0.32847 | -0.0040096 | 1e-10 | 5e-11 | 2.8868e-11 | -2.1654e-07 | 91.699 | 183.3981 | 91.6998 | 183.3996 | 183.4 | 5.835 | OK |  | 0.04517 | 183.4 | 0.033787 | 5.9951
220 | 0.1209 |  |  |  | 0.003036 | 1.5399e-05 | 196.9234 | 0.50781 | 0.025554 | 0.02564 | 0.025596 | 0.17 | 0.0006417 | 1.1351e-06 | 507.9181 | 0.19688 | 0.19688 | 2.29 | 1.145 | 0.0027686 | 0.0013843 | 0.21137 | 0.00029259 | 0.24201 | -0.013697 | 0.048228 | 0.15336 | 0.050135 | 0.041 | 0.01 | 0.005 | 0.02 | 0.050304 | 0.21696 | 71.4616 | 0.05 | -0.125 | 0.02 | 0.02 | 0.039526 | 0.019763 | 3.0911e-06 | 1.7846e-06 | 0.019763 | 0.086 | 0 | 0.00153 | 75.7667 | 0.12292 | 1e-10 | 5e-11 | 2.8868e-11 | 2.3042e-07 | 0.01 | 0.0057735 | 0.124% | -0.1 | -0.17 | 0.016% | 0.009238% | 0.02 | 0.187 | 0.093501 | 3.0193e-06 | 1.7432e-06 | 0.093501 | -0.0088176 | 1e-10 | 5e-11 | 2.8868e-11 | -2.2424e-07 | 75.776 | 151.5519 | 75.7762 | 151.5524 | 151.6 | 0.16551 | OK |  | 0.03874 | 151.6 | 0.0010935 | 0.33522
230 | 0.127 |  |  |  | 0.022015 | 0.00014225 | 154.7615 | 0.64616 | 0.026466 | 0.026 | 0.025956 | 0.17 | 0.0045878 | 1.5588e-05 | 294.1921 | 0.33991 | 0.33991 | 2.02 | 1.01 | 0.0025654 | 0.0012827 | 0.20839 | 0.0002673 | 0.21048 | -0.015623 | 0.046368 | 0.14745 | 0.048929 | 0.041 | 0.01 | 0.005 | 0.02 | 0.063507 | 0.21696 | 9.8549 | 0.05 | -0.125 | 0.02 | 0.02 | 0.0054508 | 0.0027254 | 3.6605e-06 | 2.1134e-06 | 0.0027254 | 0.086 | 0 | 0.00153 | 30.1183 | 0.45154 | 1e-10 | 5e-11 | 2.8868e-11 | 9.1633e-08 | 0.005 | 0.0028868 | 0.124% | -0.1 | -0.17 | 0.017% | 0.009815% | 0.02 | 0.026157 | 0.013078 | 3.1376e-06 | 1.8115e-06 | 0.013078 | -0.17459 | 1e-11 | 5e-12 | 2.8868e-12 | -3.2333e-08 | 30.1399 | 60.2798 | 30.1409 | 60.2818 | 60.3 | 1.9647 | OK |  | 0.015959 | 60.3 | 0.031953 | 1.7914
240 | 0.1304 |  |  |  | 0.10353 | 0.00073212 | 141.4131 | 0.70715 | 0.027128 | 0.02677 | 0.026724 | 0.17 | 0.021538 | 6.1994e-05 | 347.407 | 0.28785 | 0.28785 | 2 | 1 | 0.002608 | 0.001304 | 0.20803 | 0.00027128 | 0.20803 | -0.017969 | 0.044546 | 0.14166 | 0.048034 | 0.041 | 0.01 | 0.005 | 0.02 | 0.034274 | 0.21696 | 2.0956 | 0.05 | -0.125 | 0.02 | 0.02 | 0.0011591 | 0.00057954 | 6.1059e-06 | 3.5253e-06 | 0.00057955 | 0.086 | 0 | 0.00153 | 8.3921 | 0.6478 | 1e-09 | 5e-10 | 2.8868e-10 | 2.5543e-07 | 0.019 | 0.01097 | 0.124% | -0.1 | -0.17 | 0.017% | 0.009815% | 0.02 | 0.0055716 | 0.0027858 | 3.6461e-06 | 2.1051e-06 | 0.0027858 | 0.77294 | 1e-11 | 5e-12 | 2.8868e-12 | 3.5992e-08 | 8.4808 | 16.9615 | 8.5176 | 17.0353 | 17.1 | 1.508 | OK |  | 0.0046388 | 17.1 | 0.08682 | 1.3355
250 | 0.1219 |  |  |  | 0.26965 | 0.0016792 | 160.5841 | 0.62273 | 0.025184 | 0.02473 | 0.024692 | 0.153 | 0.055708 | 9.2275e-05 | 603.7025 | 0.16564 | 0.16564 | 1.64 | 0.82 | 0.0019992 | 0.00099958 | 0.20659 | 0.00020651 | 0.16941 | -0.02086 | 0.042763 | 0.13598 | 0.047579 | 0.041 | 0.01 | 0.005 | 0.02 | 0.14818 | 0.21696 | 0.80458 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00044502 | 0.00022251 | 1.109e-05 | 6.4026e-06 | 0.0002226 | 0.086 | 0 | 0.00153 | 3.3969 | 0.60157 | 1e-09 | 5e-10 | 2.8868e-10 | 1.0342e-07 | 0.003 | 0.0017321 | 0.124% | -0.09 | -0.153 | 0.015% | 0.00866% | 0.02 | 0.0021541 | 0.001077 | 4.6712e-06 | 2.6969e-06 | 0.001077 | 0.21872 | 1e-10 | 5e-11 | 2.8868e-11 | 6.8423e-08 | 3.5576 | 7.1153 | 3.5677 | 7.1354 | 7.2 | 1.9901 | OK |  | 0.0018132 | 7.2 | 0.27009 | 1.8341
260 | 0.1103 |  |  |  | 0.52013 | 0.002895 | 179.6616 | 0.5566 | 0.022741 | 0.022396 | 0.022364 | 0.1445 | 0.10724 | 0.00038182 | 280.8568 | 0.35605 | 0.35605 | 1.66 | 0.83 | 0.001831 | 0.00091549 | 0.20617 | 0.00018875 | 0.17113 | -0.024469 | 0.041016 | 0.13043 | 0.04776 | 0.041 | 0.01 | 0.005 | 0.02 | 0.17063 | 0.21696 | 0.41712 | 0.02 | -0.05 | 0.02 | 0.02 | 0.0011536 | 0.00057678 | 1.6403e-05 | 9.47e-06 | 0.00057686 | 0.086 | 0 | 0.00153 | 1.7899 | 0.54664 | 1e-09 | 5e-10 | 2.8868e-10 | 5.4508e-08 | 0.009 | 0.0051962 | 0.124% | -0.085 | -0.1445 | 0.014% | 0.008083% | 0.02 | 0.001119 | 0.00055951 | 6.2171e-06 | 3.5894e-06 | 0.00055952 | 0.40742 | 1e-10 | 5e-11 | 2.8868e-11 | 3.0803e-08 | 2.0694 | 4.1387 | 2.1141 | 4.2283 | 4.3 | 1.6877 | OK | Ask for accreditation in this range | 0.00097787 | 4.3 | 0.38406 | 1.5408
270 | 0.1001 |  |  |  | 0.8587 | 0.0046691 | 183.9103 | 0.54374 | 0.02064 | 0.02025 | 0.020222 | 0.136 | 0.17706 | 0.000586 | 302.1564 | 0.33095 | 0.33095 | 1.64 | 0.82 | 0.0016416 | 0.00082082 | 0.2062 | 0.00016925 | 0.16908 | -0.029035 | 0.039308 | 0.125 | 0.048869 | 0.041 | 0.01 | 0.005 | 0.02 | 0.057581 | 0.21696 | 0.25266 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00069873 | 0.00034936 | 2.3174e-05 | 1.338e-05 | 0.00034962 | 0.086 | 0 | 0.00153 | 1.0917 | 0.53781 | 1e-09 | 5e-10 | 2.8868e-10 | 3.3251e-08 | 0.015 | 0.0086603 | 0.124% | -0.08 | -0.136 | 0.012% | 0.006928% | 0.02 | 0.00067772 | 0.00033886 | 8.3119e-06 | 4.7989e-06 | 0.0003389 | 0.35832 | 1e-09 | 5e-10 | 2.8868e-10 | 1.7651e-07 | 1.4895 | 2.979 | 1.5382 | 3.0764 | 3.1 | 2.0672 | OK |  | 0.00063986 | 3.1 | 0.64549 | 1.9283
280 | 0.1025 |  |  |  | 1.2699 | 0.013756 | 92.319 | 1.0832 | 0.021977 | 0.02146 | 0.02142 | 0.187 | 0.27228 | 0.0014241 | 191.1934 | 0.52303 | 0.52303 | 1.63 | 0.815 | 0.0016707 | 0.00083537 | 0.21441 | 0.00017911 | 0.17474 | -0.034897 | 0.037638 | 0.11969 | 0.051326 | 0.041 | 0.01 | 0.005 | 0.02 | 0.2365 | 0.21696 | 0.17084 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00047247 | 0.00023624 | 3.1398e-05 | 1.8128e-05 | 0.00023693 | 0.086 | 0 | 0.00153 | 0.74064 | 1.0752 | 1e-08 | 5e-09 | 2.8868e-09 | 2.2564e-07 | 0.014 | 0.0080829 | 0.124% | -0.11 | -0.187 | 0.019% | 0.01097% | 0.02 | 0.00044072 | 0.00022036 | 1.1168e-05 | 6.4481e-06 | 0.00022046 | 0.55036 | 1e-09 | 5e-10 | 2.8868e-10 | 1.1156e-07 | 1.577 | 3.154 | 1.6808 | 3.3617 | 3.4 | 2.6005 | OK |  | 0.00074721 | 3.4 | 0.73443 | 2.4086
290 | 0.1147 |  |  |  | 1.8525 | 0.018634 | 99.4169 | 1.0059 | 0.026723 | 0.02618 | 0.026122 | 0.221 | 0.4316 | 0.001489 | 289.8619 | 0.34499 | 0.34499 | 1.64 | 0.82 | 0.0018811 | 0.00094054 | 0.23298 | 0.00021913 | 0.19105 | -0.042529 | 0.036005 | 0.1145 | 0.055723 | 0.041 | 0.01 | 0.005 | 0.02 | 0.29071 | 0.21696 | 0.11712 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00032389 | 0.00016194 | 4.305e-05 | 2.4855e-05 | 0.00016384 | 0.086 | 0 | 0.00153 | 0.50885 | 1.0007 | 1e-08 | 5e-09 | 2.8868e-09 | 1.5504e-07 | 0.021 | 0.012124 | 0.124% | -0.13 | -0.221 | 0.022% | 0.0127% | 0.02 | 0.00027803 | 0.00013902 | 1.5948e-05 | 9.2076e-06 | 0.00013932 | 0.35615 | 1e-09 | 5e-10 | 2.8868e-10 | 6.9047e-08 | 1.4422 | 2.8845 | 1.5021 | 3.0042 | 3.1 | 2.3015 | OK |  | 0.00082842 | 3 | 0.71694 | 2.0754
300 | 0.1295 | -9.3e-05 | 0.12948 | 0.018603 | 2.5414 | 0.02538 | 100.1342 | 0.99866 | 0.031144 | 0.03063 | 0.030562 | 0.221 | 0.61118 | 0.0019461 | 314.063 | 0.31841 | 0.31841 | 1.63 | 0.815 | 0.0021109 | 0.0010554 | 0.24049 | 0.00025382 | 0.196 | -0.052594 | 0.03441 | 0.10942 | 0.06285 | 0.041 | 0.01 | 0.005 | 0.02 | 0.15971 | 0.21696 | 0.085369 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00023609 | 0.00011805 | 5.6828e-05 | 3.2809e-05 | 0.00012252 | 0.086 | 0 | 0.00153 | 0.37137 | 0.99495 | 1e-08 | 5e-09 | 2.8868e-09 | 1.1317e-07 | 0.002 | 0.0011547 | 0.124% | -0.13 | -0.221 | 0.021% | 0.01212% | 0.02 | 0.0009817 | 0.00049085 | 1.8224e-05 | 1.0521e-05 | 0.00049096 | 0.32561 | 1e-09 | 5e-10 | 2.8868e-10 | 4.8301e-08 | 1.3715 | 2.7431 | 1.427 | 2.854 | 2.9 | 1.9024 | OK |  | 0.00090317 | 2.9 | 0.63904 | 1.6772
310 | 0.1375 | -5.9765e-05 | 0.13748 | 0.011954 | 3.1295 | 0.03824 | 81.8391 | 1.2219 | 0.03329 | 0.032732 | 0.032671 | 0.187 | 0.7577 | 0.003172 | 238.8678 | 0.41864 | 0.41864 | 1.63 | 0.815 | 0.0022412 | 0.0011206 | 0.24211 | 0.00027132 | 0.19732 | -0.065976 | 0.032853 | 0.10447 | 0.073703 | 0.041 | 0.01 | 0.005 | 0.02 | 0.056636 | 0.21696 | 0.069325 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00019172 | 9.5861e-05 | 6.8591e-05 | 3.9601e-05 | 0.00010372 | 0.086 | 0 | 0.00153 | 0.30173 | 1.2182 | 1e-08 | 5e-09 | 2.8868e-09 | 9.1964e-08 | 0.001 | 0.00057735 | 0.124% | -0.11 | -0.187 | 0.018% | 0.01039% | 0.02 | 0.00079187 | 0.00039594 | 2.1154e-05 | 1.2213e-05 | 0.00039613 | 0.42625 | 1e-09 | 5e-10 | 2.8868e-10 | 3.8791e-08 | 1.5191 | 3.0381 | 1.5889 | 3.1778 | 3.2 | 1.8961 | OK |  | 0.0010653 | 3.2 | 0.57921 | 1.7056
320 | 0.1422 | -2.2e-05 | 0.14219 | 0.0044002 | 3.6 | 0.036751 | 97.9563 | 1.0209 | 0.034161 | 0.033631 | 0.033574 | 0.17 | 0.86484 | 0.0029152 | 296.6646 | 0.33708 | 0.33708 | 1.63 | 0.815 | 0.0023179 | 0.0011589 | 0.24024 | 0.00027842 | 0.19579 | -0.083731 | 0.031334 | 0.099641 | 0.089402 | 0.041 | 0.01 | 0.005 | 0.02 | 0.031185 | 0.21696 | 0.060266 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00016667 | 8.3334e-05 | 7.7999e-05 | 4.5033e-05 | 9.4724e-05 | 0.086 | 0 | 0.00153 | 0.26237 | 1.0182 | 1e-08 | 5e-09 | 2.8868e-09 | 7.9978e-08 | 0.001 | 0.00057735 | 0.124% | -0.1 | -0.17 | 0.016% | 0.009238% | 0.02 | 0.00069377 | 0.00034689 | 2.3297e-05 | 1.345e-05 | 0.00034715 | 0.34243 | 1e-09 | 5e-10 | 2.8868e-10 | 3.3909e-08 | 1.3558 | 2.7117 | 1.4089 | 2.8177 | 2.9 | 1.7502 | OK |  | 0.00099068 | 2.8 | 0.59036 | 1.5772
330 | 0.1455 | 2.0795e-05 | 0.14551 | 0.0041587 | 4.0234 | 0.041553 | 96.8245 | 1.0328 | 0.03493 | 0.034394 | 0.034341 | 0.153 | 0.96588 | 0.0035652 | 270.9195 | 0.36911 | 0.36911 | 1.63 | 0.815 | 0.0023716 | 0.0011858 | 0.24007 | 0.00028468 | 0.19565 | -0.1067 | 0.029852 | 0.09493 | 0.1108 | 0.041 | 0.01 | 0.005 | 0.02 | 0.035248 | 0.21696 | 0.053924 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00014913 | 7.4564e-05 | 8.6468e-05 | 4.9922e-05 | 8.9733e-05 | 0.086 | 0 | 0.00153 | 0.2348 | 1.0304 | 1e-08 | 5e-09 | 2.8868e-09 | 7.1581e-08 | 0.001 | 0.00057735 | 0.124% | -0.09 | -0.153 | 0.015% | 0.00866% | 0.02 | 0.0006212 | 0.0003106 | 2.5318e-05 | 1.4617e-05 | 0.00031094 | 0.37435 | 1e-09 | 5e-10 | 2.8868e-10 | 3.0312e-08 | 1.3617 | 2.7233 | 1.4206 | 2.8412 | 2.9 | 1.7128 | OK |  | 0.001013 | 2.9 | 0.57845 | 1.5572
340 | 0.1478 | 5.3e-05 | 0.14782 | 0.010599 | 4.414 | 0.044489 | 99.2158 | 1.0079 | 0.035789 | 0.03534 | 0.035286 | 0.153 | 1.0688 | 0.0060864 | 175.6074 | 0.56945 | 0.56945 | 1.63 | 0.815 | 0.0024091 | 0.0012046 | 0.24214 | 0.00029168 | 0.19735 | -0.13435 | 0.028408 | 0.090339 | 0.13732 | 0.041 | 0.01 | 0.005 | 0.02 | 0.037258 | 0.21696 | 0.049152 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00013593 | 6.7966e-05 | 9.428e-05 | 5.4432e-05 | 8.7076e-05 | 0.086 | 0 | 0.00153 | 0.21403 | 1.0057 | 1e-08 | 5e-09 | 2.8868e-09 | 6.526e-08 | 0.002 | 0.0011547 | 0.124% | -0.09 | -0.153 | 0.015% | 0.00866% | 0.02 | 0.00056137 | 0.00028068 | 2.7376e-05 | 1.5806e-05 | 0.00028113 | 0.57675 | 1e-07 | 5e-08 | 2.8868e-08 | 2.7355e-06 | 1.3422 | 2.6843 | 1.469 | 2.9379 | 3 | 1.4254 | OK |  | 0.0010737 | 3 | 0.46682 | 1.2702
350 | 0.1481 | 5.8586e-05 | 0.14812 | 0.011716 | 4.6851 | 0.048847 | 95.9137 | 1.0426 | 0.03638 | 0.035982 | 0.035933 | 0.136 | 1.1509 | 0.0059386 | 193.797 | 0.516 | 0.516 | 1.63 | 0.815 | 0.002414 | 0.001207 | 0.24564 | 0.0002965 | 0.2002 | -0.16239 | 0.027003 | 0.085868 | 0.16462 | 0.041 | 0.01 | 0.005 | 0.02 | 0.016219 | 0.21696 | 0.046307 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00085376 | 0.00042688 | 0.00011028 | 6.3668e-05 | 0.0004316 | 0.086 | 0 | 0.00153 | 0.20167 | 1.0405 | 1e-08 | 5e-09 | 2.8868e-09 | 6.1491e-08 | 0.005 | 0.0028868 | 0.124% | -0.08 | -0.136 | 0.013% | 0.007506% | 0.02 | 0.00052134 | 0.00026067 | 2.9018e-05 | 1.6753e-05 | 0.00026121 | 0.52214 | 1e-08 | 5e-09 | 2.8868e-09 | 2.5381e-07 | 1.3691 | 2.7382 | 1.4717 | 2.9435 | 3 | 1.2437 | OK |  | 0.0010914 | 3 | 0.40804 | 1.1061
360 | 0.1447 | 3.6e-05 | 0.14471 | 0.0071995 | 4.7681 | 0.049975 | 95.4099 | 1.0481 | 0.036439 | 0.03585 | 0.03581 | 0.1105 | 1.2007 | 0.0058361 | 205.7428 | 0.48604 | 0.48604 | 1.63 | 0.815 | 0.0023586 | 0.0011793 | 0.25183 | 0.00029698 | 0.20524 | -0.18132 | 0.025634 | 0.081517 | 0.18312 | 0.041 | 0.01 | 0.005 | 0.02 | 0.090397 | 0.21696 | 0.045501 | 0.02 | -0.05 | 0.02 | 0.02 | 0.0008389 | 0.00041945 | 0.00011152 | 6.4387e-05 | 0.00042436 | 0.086 | 0 | 0.00153 | 0.19815 | 1.046 | 1e-08 | 5e-09 | 2.8868e-09 | 6.0423e-08 | 0.013 | 0.0075056 | 0.124% | -0.065 | -0.1105 | 0.01% | 0.005774% | 0.02 | 0.00049969 | 0.00024984 | 3.0015e-05 | 1.7329e-05 | 0.00025045 | 0.49158 | 1e-08 | 5e-09 | 2.8868e-09 | 2.4315e-07 | 1.378 | 2.7561 | 1.4674 | 2.9348 | 3 | 1.7565 | OK |  | 0.0010932 | 2.9 | 0.57301 | 1.6441
370 | 0.1449 | -1.1015e-05 | 0.1449 | 0.0022031 | 4.9339 | 0.053122 | 92.8796 | 1.0767 | 0.042312 | 0.04117 | 0.041065 | 0.255 | 1.4407 | 0.0090548 | 159.1143 | 0.62848 | 0.62848 | 1.33 | 0.665 | 0.0019272 | 0.00096358 | 0.29201 | 0.00028137 | 0.19419 | -0.18132 | 0.024304 | 0.077287 | 0.18294 | 0.041 | 0.01 | 0.005 | 0.02 | 0.46638 | 0.21696 | 0.043972 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00081071 | 0.00040536 | 0.00011401 | 6.5823e-05 | 0.00041067 | 0.086 | 0 | 0.00153 | 0.19149 | 1.0746 | 1e-08 | 5e-09 | 2.8868e-09 | 5.8396e-08 | 0.013 | 0.0075056 | 0.124% | -0.15 | -0.255 | 0.024% | 0.01386% | 0.02 | 0.00041645 | 0.00020823 | 3.4815e-05 | 2.01e-05 | 0.00020919 | 0.63443 | 1e-08 | 5e-09 | 2.8868e-09 | 2.0226e-07 | 1.3943 | 2.7887 | 1.5531 | 3.1063 | 3.2 | 3.0366 | OK |  | 0.001354 | 3.1 | 0.90445 | 2.7739
380 | 0.1567 | -7.2e-05 | 0.15668 | 0.014402 | 5.47 | 0.057237 | 95.5676 | 1.0464 | 0.061621 | 0.0604 | 0.060195 | 0.34 | 2.151 | 0.014545 | 147.884 | 0.67621 | 0.67621 | 0.79 | 0.395 | 0.0012379 | 0.00061897 | 0.39324 | 0.0002434 | 0.15533 | -0.16239 | 0.023012 | 0.073177 | 0.16401 | 0.041 | 0.01 | 0.005 | 0.02 | 0.52948 | 0.21696 | 0.039663 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00073126 | 0.00036563 | 0.00012205 | 7.0466e-05 | 0.00037236 | 0.086 | 0 | 0.00153 | 0.17275 | 1.0446 | 1e-08 | 5e-09 | 2.8868e-09 | 5.2683e-08 | 0.004 | 0.0023094 | 0.124% | -0.2 | -0.34 | 0.032% | 0.01848% | 0.02 | 0.00027894 | 0.00013947 | 4.9021e-05 | 2.8302e-05 | 0.00014231 | 0.68048 | 1e-08 | 5e-09 | 2.8868e-09 | 1.3505e-07 | 1.2821 | 2.5642 | 1.4909 | 2.9819 | 3 | 2.3691 | OK |  | 0.0018486 | 3 | 0.75225 | 2.021
390 | 0.1691 | -0.0001339 | 0.16905 | 0.026787 | 6.0314 | 0.068257 | 88.3631 | 1.1317 | 0.081511 | 0.08051 | 0.080305 | 0.255 | 2.9073 | 0.01298 | 223.9843 | 0.44646 | 0.44646 | 0.65 | 0.325 | 0.0010991 | 0.00054957 | 0.48203 | 0.00026491 | 0.15666 | -0.13435 | 0.021757 | 0.069187 | 0.1361 | 0.041 | 0.01 | 0.005 | 0.02 | 0.32073 | 0.21696 | 0.035971 | 0.02 | -0.05 | 0.02 | 0.02 | 0.0006632 | 0.0003316 | 0.00013047 | 7.5327e-05 | 0.00034005 | 0.086 | 0 | 0.00153 | 0.15668 | 1.1299 | 1e-08 | 5e-09 | 2.8868e-09 | 4.7787e-08 | 0.002 | 0.0011547 | 0.124% | -0.15 | -0.255 | 0.025% | 0.01443% | 0.02 | 0.00020638 | 0.00010319 | 6.4146e-05 | 3.7035e-05 | 0.00010963 | 0.44855 | 1e-08 | 5e-09 | 2.8868e-09 | 9.9757e-08 | 1.2601 | 2.5202 | 1.3618 | 2.7236 | 2.8 | 1.5022 | OK |  | 0.0022823 | 2.7 | 0.52211 | 1.2434
400 | 0.1798 | -0.00018 | 0.17974 | 0.036013 | 6.6076 | 0.070879 | 93.2238 | 1.0727 | 0.096899 | 0.09599 | 0.095802 | 0.1955 | 3.561 | 0.017099 | 208.2516 | 0.48019 | 0.48019 | 0.49 | 0.245 | 0.00088102 | 0.00044051 | 0.53892 | 0.0002374 | 0.13204 | -0.1067 | 0.02054 | 0.065317 | 0.10866 | 0.041 | 0.01 | 0.005 | 0.02 | 0.22242 | 0.21696 | 0.032834 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00060536 | 0.00030268 | 0.00013911 | 8.0317e-05 | 0.00031316 | 0.086 | 0 | 0.00153 | 0.14303 | 1.0712 | 1e-08 | 5e-09 | 2.8868e-09 | 4.3626e-08 | 0.001 | 0.00057735 | 0.124% | -0.115 | -0.1955 | 0.019% | 0.01097% | 0.02 | 0.00016849 | 8.4246e-05 | 7.722e-05 | 4.4583e-05 | 9.5316e-05 | 0.48202 | 1e-08 | 5e-09 | 2.8868e-09 | 8.1375e-08 | 1.1613 | 2.3227 | 1.2727 | 2.5453 | 2.6 | 1.1442 | OK |  | 0.0025194 | 2.6 | 0.43165 | 0.94651
0.1105
0.34
[46 empty] | with filter | 4e-10
[46 empty] | without filter | 4.8e-05
0.00083333
Note on CF20: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: En values will be very small, because Uc is very large

## Temp coeff
CCPR-k2.b 2004 report |  | CCPR-k2.c 2003 report (@ 25 °C)
Si Photodiode |  | PtSi
nm | [1/K] | nm | %/°C
300 | -9.3e-05 | 200 | -0.09
320 | -2.2e-05 | 210 | -0.1
340 | 5.3e-05 | 220 | -0.1
360 | 3.6e-05 | 230 | -0.1
380 | -7.2e-05 | 240 | -0.1
400 | -0.00018 | 250 | -0.09
450 | -0.00011 | 260 | -0.085
500 | -8.1e-05 | 270 | -0.08
550 | -7.1e-05 | 280 | -0.11
600 | -6.1e-05 | 290 | -0.13
650 | -5.1e-05 | 300 | -0.13
700 | -4.9e-05 | 310 | -0.11
750 | -4.6e-05 | 320 | -0.1
800 | -4.4e-05 | 330 | -0.09
850 | -4.2e-05 | 340 | -0.09
900 | -4.4e-05 | 350 | -0.08
950 | 8.9e-05 | 360 | -0.065
1000 | 0.0018 | 370 | -0.15
|  | 380 | -0.2
|  | 390 | -0.15
|  | 400 | -0.115

## Wavelength Uc
Wavelength [nm] | Calculated response [A/W] |  |  |  |  | sensitivity coefficient | Δλ | coverage factor, k | Ucλ |  | (Ucλ)2 | AVE |  |  | sensitivity coefficient | Ucλ |  | (Ucλ)2 | worst case using ration of signals
|  | +0,3 nm | [A/W] | -0,3 nm | [A/W] | [1/nm] | [nm] | [1] | [%] |  | [%] | signals [V/V] |  |  | [1/nm] | [%] |  | [%]
λ0 | sT0 | λ+ | sT+ | λ- | sT- | δsT/δλ | 0.3 | 2 | Using response of UUT |  |  | IT0/IS0 | (IT0/IS0)+ | (IT0/IS0)- | δ(IT/IS)/δλ | using ratio of signals | [5 empty] | Max Uc [%]
200 | 0.023665 | 200.3 | 0.023694 | 199.7 | 0.023636 | 0.0040792 |  |  | 0.061189 | 0.061189 | 0.003744 | 0.20869 | 0.2088 | 0.20858 | 0.0018102 | 0.027152 | 0.027152 | 0.00073726 | 0.0030068 | 1 | 19 | 0.061189
210 | 0.024629 | 210.3 | 0.024658 | 209.7 | 0.0246 | 0.0039014 |  |  | 0.058521 | 0.058521 | 0.0034247 | 0.21177 | 0.21182 | 0.21172 | 0.00080063 | 0.012009 | 0.012009 | 0.00014423 | 0.0032805 | 1 | 2 | 0.058521
220 | 0.025554 | 220.3 | 0.02558 | 219.7 | 0.025528 | 0.0033536 |  |  | 0.050304 | 0.050304 | 0.0025305 | 0.21137 | 0.21129 | 0.21144 | -0.001199 | -0.017985 | 0.017985 | 0.00032348 | 0.002207 | 1 |  | 0.050304
230 | 0.026466 | 230.3 | 0.0265 | 229.7 | 0.026432 | 0.0042338 |  |  | 0.063507 | 0.063507 | 0.0040332 | 0.20839 | 0.20834 | 0.20845 | -0.00081642 | -0.012246 | 0.012246 | 0.00014997 | 0.0038832 | 1 |  | 0.063507
240 | 0.027128 | 240.3 | 0.027107 | 239.7 | 0.027144 | -0.0022849 |  |  | -0.034274 | 0.034274 | 0.0011747 | 0.20803 | 0.20801 | 0.20805 | -0.00031539 | -0.0047309 | 0.0047309 | 2.2381e-05 | 0.0011523 | 1 |  | 0.034274
250 | 0.025184 | 250.3 | 0.025109 | 249.7 | 0.025258 | -0.0098787 |  |  | -0.14818 | 0.14818 | 0.021958 | 0.20659 | 0.20656 | 0.20663 | -0.00051991 | -0.0077986 | 0.0077986 | 6.0818e-05 | 0.021897 | 1 |  | 0.14818
260 | 0.022741 | 260.3 | 0.022663 | 259.7 | 0.022818 | -0.011375 |  |  | -0.17063 | 0.17063 | 0.029114 | 0.20617 | 0.20616 | 0.20619 | -0.00030144 | -0.0045217 | 0.0045217 | 2.0446e-05 | 0.029093 | 1 |  | 0.17063
270 | 0.02064 | 270.3 | 0.020619 | 269.7 | 0.020666 | -0.0038387 |  |  | -0.057581 | 0.057581 | 0.0033156 | 0.2062 | 0.20627 | 0.20613 | 0.0011558 | 0.017337 | 0.017337 | 0.00030057 | 0.003015 | 1 |  | 0.057581
280 | 0.021977 | 280.3 | 0.022083 | 279.7 | 0.021875 | 0.015767 |  |  | 0.2365 | 0.2365 | 0.055931 | 0.21441 | 0.21489 | 0.21394 | 0.0073667 | 0.1105 | 0.1105 | 0.01221 | 0.043721 | 1 |  | 0.2365
290 | 0.026723 | 290.3 | 0.026878 | 289.7 | 0.026568 | 0.019381 |  |  | 0.29071 | 0.29071 | 0.084515 | 0.23298 | 0.23342 | 0.23253 | 0.0063467 | 0.095201 | 0.095201 | 0.0090632 | 0.075452 | 1 |  | 0.29071
300 | 0.031144 | 300.3 | 0.031242 | 299.7 | 0.031043 | 0.010648 |  |  | 0.15971 | 0.15971 | 0.025508 | 0.24049 | 0.24059 | 0.24039 | 0.0013725 | 0.020588 | 0.020588 | 0.00042386 | 0.025084 | 1 |  | 0.15971
310 | 0.03329 | 310.3 | 0.033327 | 309.7 | 0.033252 | 0.0037757 |  |  | 0.056636 | 0.056636 | 0.0032077 | 0.24211 | 0.24209 | 0.24213 | -0.00025102 | -0.0037654 | 0.0037654 | 1.4178e-05 | 0.0031935 | 1 |  | 0.056636
320 | 0.034161 | 320.3 | 0.034183 | 319.7 | 0.03414 | 0.002079 |  |  | 0.031185 | 0.031185 | 0.00097247 | 0.24024 | 0.24019 | 0.24029 | -0.00068165 | -0.010225 | 0.010225 | 0.00010455 | 0.00086793 | 1 |  | 0.031185
330 | 0.03493 | 330.3 | 0.034954 | 329.7 | 0.034905 | 0.0023498 |  |  | 0.035248 | 0.035248 | 0.0012424 | 0.24007 | 0.2401 | 0.24004 | 0.00042703 | 0.0064054 | 0.0064054 | 4.1029e-05 | 0.0012014 | 1 |  | 0.035248
340 | 0.035789 | 340.3 | 0.035816 | 339.7 | 0.035762 | 0.0024839 |  |  | 0.037258 | 0.037258 | 0.0013882 | 0.24214 | 0.24224 | 0.24205 | 0.0013479 | 0.020219 | 0.020219 | 0.0004088 | 0.00097939 | 1 |  | 0.037258
350 | 0.03638 | 350.3 | 0.036378 | 349.7 | 0.03638 | -7.066e-05 |  |  | -0.0010599 | 0.0010599 | 1.1234e-06 | 0.24564 | 0.24572 | 0.24556 | 0.0010812 | 0.016219 | 0.016219 | 0.00026304 | -0.00026192 |  |  | 0.016219
360 | 0.036439 | 360.3 | 0.036476 | 359.7 | 0.036406 | 0.0032057 |  |  | 0.048085 | 0.048085 | 0.0023122 | 0.25183 | 0.2523 | 0.25139 | 0.0060264 | 0.090397 | 0.090397 | 0.0081715 | -0.0058594 |  |  | 0.090397
370 | 0.042312 | 370.3 | 0.042715 | 369.7 | 0.041926 | 0.031092 |  |  | 0.46638 | 0.46638 | 0.21751 | 0.29201 | 0.29432 | 0.28978 | 0.025943 | 0.38915 | 0.38915 | 0.15144 | 0.066071 | 1 |  | 0.46638
380 | 0.061621 | 380.3 | 0.062272 | 379.7 | 0.060967 | 0.035299 |  |  | 0.52948 | 0.52948 | 0.28035 | 0.39324 | 0.39641 | 0.39005 | 0.02696 | 0.4044 | 0.4044 | 0.16354 | 0.11681 | 1 |  | 0.52948
390 | 0.081511 | 390.3 | 0.082031 | 389.7 | 0.080985 | 0.021382 |  |  | 0.32073 | 0.32073 | 0.10287 | 0.48203 | 0.48412 | 0.4799 | 0.014563 | 0.21844 | 0.21844 | 0.047718 | 0.055148 | 1 |  | 0.32073
400 | 0.096899 | 400.3 | 0.09733 | 399.7 | 0.096468 | 0.014828 |  |  | 0.22242 | 0.22242 | 0.04947 | 0.53892 | 0.54043 | 0.53742 | 0.0093244 | 0.13987 | 0.13987 | 0.019563 | 0.029907 | 1 |  | 0.22242

## Source stability
lamp stability (short-term)
Xe source, Intensity controller with monitor detector, measured for 2 min with 1s interval.
(Xenon source)
average [V] | 0.23603
std dev [V] | 6.702e-05
min [V] | 0.2359
max [V] | 0.23618
std dev [%] | 0.028394
dev between min & max [%] | 0.11855
lamp drift (long-term)
Xe source, Intensity controller with monitor detector, measured for 1 hour with 5s interval.
(Xenon source)
average [V] | 0.23614
std dev [V] | 8.0325e-05
min [V] | 0.23594
max [V] | 0.23645
std dev [%] | 0.034016
dev between min & max [%] | 0.21696
0.034016

## DVM
DVM Voltage | ORE-556.RA
10 V range
measured value [V] | 3.1295
DVM Uc on cert [V] | 6e-06
Uc [%] k = 2 | 0.00019172
Uc [%] k = 1 | 9.5861e-05
Manufacturer's spec
Accuracy | ± (0,002 % of reading + 0,0006 % of range)
reading [V] | 3.1295
range [V] | 1
10 ppm = 0,001 %
% of reading | 0.002
% of range | 0.0006
Accuracy (24 h) [%] | 6.8591e-05
Divisor sqrt(3) | 3.9601e-05 | 0.00010372
Sqrt(sumsq) of Uc + Accuracy | 0.00010372 | 0.00034962
DVM Voltage | ORE-556.RA
1 V range
measured value [V] | 3.1295
DVM Uc on cert [V] | 6e-06
Uc [%] k = 2 | 0.00019172
Uc [%] k = 1 | 9.5861e-05
Manufacturer's spec
Accuracy | ± (0,002 % of reading + 0,0006 % of range)
reading [V] | 3.1295
range [V] | 1
10 ppm = 0,001 %
% of reading | 0.002
% of range | 0.0006
Accuracy (24 h) [%] | 6.8591e-05
Divisor sqrt(3) | 3.9601e-05 | 0.00010372
Sqrt(sumsq) of Uc + Accuracy | 0.00010372 | 0.00034962
DVM Voltage | ORE-556.RA
100 mV range
measured value [V] | 0.45403
DVM Uc on cert [V] | 1.2e-06
Uc [%] | 0.0002643
Manufacturer's spec
Accuracy | ± (0,002 % of reading + 0,0006 % of range)
reading [V] | 0.45403
range [V] | 0.1
10 ppm = 0,001 %
% of reading | 0.003
% of range | 0.003
Accuracy (24 h) [%] | 1.6621e-05
Sqrt(sumsq) of Uc + Accuracy | 0.00026482

## Drift is STD
Predicting wavlength dependent responsivity-drift is silicon, Littler et al (NewRad proceedings)
λ [nm] | Drift in response (per annum) [%] | Std dev of drift rate σ(λ) (Prediction contribution U per annum) [%] | 3.18 | + | - | Last calibration: 2 November 2018 | PtSi calibrated on 18 March 2019 | Time since last calibration: Max 6 month
200 | -0.01076 | 0.05206 | 0.16555 | 0.15479 | -0.17631
210 | -0.012099 | 0.050125 | 0.1594 | 0.1473 | -0.1715
220 | -0.013697 | 0.048228 | 0.15336 | 0.13967 | -0.16706
230 | -0.015623 | 0.046368 | 0.14745 | 0.13183 | -0.16307
240 | -0.017969 | 0.044546 | 0.14166 | 0.12369 | -0.15963
250 | -0.02086 | 0.042763 | 0.13598 | 0.11513 | -0.15684
260 | -0.024469 | 0.041016 | 0.13043 | 0.10596 | -0.1549
270 | -0.029035 | 0.039308 | 0.125 | 0.095964 | -0.15404
280 | -0.034897 | 0.037638 | 0.11969 | 0.08479 | -0.15458
290 | -0.042529 | 0.036005 | 0.1145 | 0.071966 | -0.15702
300 | -0.052594 | 0.03441 | 0.10942 | 0.05683 | -0.16202
310 | -0.065976 | 0.032853 | 0.10447 | 0.038496 | -0.17045
320 | -0.083731 | 0.031334 | 0.099641 | 0.01591 | -0.18337
330 | -0.1067 | 0.029852 | 0.09493 | -0.011773 | -0.20163
340 | -0.13435 | 0.028408 | 0.090339 | -0.044007 | -0.22468
350 | -0.16239 | 0.027003 | 0.085868 | -0.076525 | -0.24826
360 | -0.18132 | 0.025634 | 0.081517 | -0.099802 | -0.26284
370 | -0.18132 | 0.024304 | 0.077287 | -0.10403 | -0.25861
380 | -0.16239 | 0.023012 | 0.073177 | -0.089216 | -0.23557
390 | -0.13435 | 0.021757 | 0.069187 | -0.065158 | -0.20353
400 | -0.1067 | 0.02054 | 0.065317 | -0.041385 | -0.17202
351 | -0.16489 | 0.026864 | 0.085427 | -0.079461 | -0.25032
476.486 | -0.022035 | 0.012483 | 0.039695 | 0.017659 | -0.06173
487.986 | -0.018502 | 0.011462 | 0.036451 | 0.017948 | -0.054953
514.531 | -0.012937 | 0.0092984 | 0.029569 | 0.016632 | -0.042506

## RAW
λ [Å] | Steps | Step interval | Environmental conditions |  | STD - ORW-005.RA | [10 empty] | UUT - PtSi | [6 empty] | STD responsivity [A/W] | STD Uc [%] | UUT responsivity [A/W] | PtSi cert response [A/W] | PtSi cert Uc [A/W] | PtSi cert Uc [%] | Deviation between PtSi cal cert and now [%] | Combined Uc |  | Remarks | Calculated uncertainty for this calibration [%]
[34 empty] | Gain: | 1000000 | PtSi cert response, corrected for dev from cal temp [A/W]
|  |  | Environmental conditions |  | STD - ORW-005.RA | [10 empty] | UUT - PtSi
|  |  | Temp | Humidity | Light |  |  |  |  | Dark | [5 empty] | Light |  |  | Dark
[5 empty] | Average [V] |  |  | Std Dev [V] | Std Dev [%] | Average [V] |  |  |  | Std Dev [V] | Net [V] | Average [V] | Std Dev [V] | Std Dev [%] | Average [V] |  | Std Dev [V] | Net [V] | [12 empty] | 1.7
Filter no. 1: Blank Xenon arc source |  |  | 0.2 | 1.6 | [8 empty] | noise | [6 empty] | noise | [14 empty] | %/°C |  | Beam power [W] | Beam power [uW]
2000 | -279093 | 28804.7 | 1 |  | -0.0099425 | 0.0099425 | 0.0015026 | 1.1497e-06 | -0.011563 | -0.0095037 | 0.0095037 | 0.0014362 | 95.5867 | 1.1003e-06 | -0.0004388 | -0.013421 | 4.6328e-07 | -0.0034518 | -0.013513 | 100.6823 | 6.3059e-07 | 9.1572e-05 | 0.1134 | 6.12 | 0.023665 | 0.02676 | 0.00029 | 1.0837 | 11.4287 | 6.2152 | 0.0014709 | 18/03/2019 | 191.3 | 0.045272 | -0.09 | 0.026719 | 3.8695e-09 | 0.0038695
2100 | -250288.3 | 28804.7 | 23.6 | 52.2 | -0.01036 | 0.01036 | 0.0015656 | 2.1869e-06 | -0.02111 | -0.0094971 | 0.0094971 | 0.0014352 | 91.674 | 5.7128e-07 | -0.00086255 | -0.013331 | 5.3452e-07 | -0.0040096 | -0.013514 | 101.3702 | 5.458e-07 | 0.00018267 | 0.1163 | 4.25 | 0.024629 | 0.0262 | 0.0002 | 0.76336 | 5.835 | 4.318 | 0.0010635 |  | 183.4 | 0.04517 | -0.1 | 0.026155 | 7.4166e-09 | 0.0074166
2200 | -221483.6 | 28804.7 |  |  | -0.012528 | 0.012528 | 0.0018933 | 1.5399e-05 | -0.12292 | -0.0094922 | 0.0094922 | 0.0014345 | 75.7667 | 7.4476e-07 | -0.003036 | -0.012873 | 1.1351e-06 | -0.0088176 | -0.013515 | 104.9848 | 5.5469e-07 | 0.0006417 | 0.1209 | 2.29 | 0.025554 | 0.02564 | 0.00017 | 0.66303 | 0.16551 | 2.3841 | 0.00060922 |  | 151.6 | 0.03874 | -0.1 | 0.025596 | 2.5112e-08 | 0.025112
2300 | -192678.9 | 28804.7 |  |  | -0.031503 | 0.031503 | 0.0047609 | 0.00014225 | -0.45154 | -0.0094883 | 0.0094883 | 0.0014339 | 30.1183 | 7.6356e-07 | -0.022015 | -0.0089283 | 1.5588e-05 | -0.17459 | -0.013516 | 151.3844 | 4.4865e-07 | 0.0045878 | 0.127 | 2.02 | 0.026466 | 0.026 | 0.00016 | 0.61538 | 1.9647 | 2.1117 | 0.00055887 |  | 60.3 | 0.015959 | -0.1 | 0.025956 | 1.7335e-07 | 0.17335
2400 | -163874.2 | 28804.7 |  |  | -0.11302 | 0.11302 | 0.017079 | 0.00073212 | -0.6478 | -0.0094843 | 0.0094843 | 0.0014333 | 8.3921 | 5.8971e-07 | -0.10353 | 0.0080205 | 6.1994e-05 | 0.77294 | -0.013517 | -168.5333 | 4.5523e-07 | 0.021538 | 0.1304 | 2 | 0.027128 | 0.02677 | 0.00017 | 0.63504 | 1.508 | 2.0984 | 0.00056924 |  | 17.1 | 0.0046388 | -0.1 | 0.026724 | 7.9395e-07 | 0.79395
2500 | -135069.5 | 28804.7 |  |  | -0.27913 | 0.27913 | 0.042184 | 0.0016792 | -0.60157 | -0.0094818 | 0.0094818 | 0.0014329 | 3.3969 | 6.5991e-07 | -0.26965 | 0.04219 | 9.2275e-05 | 0.21872 | -0.013519 | -32.0423 | 6.246e-07 | 0.055708 | 0.1219 | 1.64 | 0.025184 | 0.02473 | 0.00015 | 0.60655 | 1.9901 | 1.7486 | 0.00044035 |  | 7.2 | 0.0018132 | -0.09 | 0.024692 | 2.2121e-06 | 2.2121
2600 | -106264.8 | 28804.7 |  |  | -0.5296 | 0.5296 | 0.080036 | 0.002895 | -0.54664 | -0.0094791 | 0.0094791 | 0.0014325 | 1.7899 | 4.3578e-07 | -0.52013 | 0.093717 | 0.00038182 | 0.40742 | -0.013519 | -14.4258 | 5.9308e-07 | 0.10724 | 0.1103 | 1.66 | 0.022741 | 0.022396 | 9.8e-05 | 0.43758 | 1.6877 | 1.7167 | 0.0003904 |  | 4.3 | 0.00097787 | -0.085 | 0.022364 | 4.7155e-06 | 4.7155
2700 | -77460.1 | 28804.7 |  |  | -0.86818 | 0.86818 | 0.1312 | 0.0046691 | -0.53781 | -0.0094776 | 0.0094776 | 0.0014323 | 1.0917 | 6.1721e-07 | -0.8587 | 0.16354 | 0.000586 | 0.35832 | -0.013521 | -8.2673 | 5.5695e-07 | 0.17706 | 0.1001 | 1.64 | 0.02064 | 0.02025 | 0.0001 | 0.49383 | 2.0672 | 1.7127 | 0.00035352 |  | 3.1 | 0.00063986 | -0.08 | 0.020222 | 8.5784e-06 | 8.5784
2800 | -48655.4 | 28804.7 |  |  | -1.2794 | 1.2794 | 0.19335 | 0.013756 | -1.0752 | -0.0094757 | 0.0094757 | 0.001432 | 0.74064 | 5.5096e-07 | -1.2699 | 0.25876 | 0.0014241 | 0.55036 | -0.01352 | -5.2249 | 4.6737e-07 | 0.27228 | 0.1025 | 1.63 | 0.021977 | 0.02146 | 0.00013 | 0.60578 | 2.6005 | 1.7389 | 0.00038216 |  | 3.4 | 0.00074721 | -0.11 | 0.02142 | 1.2389e-05 | 12.3894
2900 | -19850.7 | 28804.7 |  |  | -1.862 | 1.862 | 0.28139 | 0.018634 | -1.0007 | -0.0094746 | 0.0094746 | 0.0014318 | 0.50885 | 3.5741e-07 | -1.8525 | 0.41808 | 0.001489 | 0.35615 | -0.01352 | -3.2338 | 4.2849e-07 | 0.4316 | 0.1147 | 1.64 | 0.026723 | 0.02618 | 0.00013 | 0.49656 | 2.3015 | 1.7135 | 0.00045791 |  | 3.1 | 0.00082842 | -0.13 | 0.026122 | 1.6151e-05 | 16.1508
3000 | 8954 | 28804.7 |  |  | -2.5509 | 2.5509 | 0.3855 | 0.02538 | -0.99495 | -0.0094731 | 0.0094731 | 0.0014316 | 0.37137 | 6.0211e-07 | -2.5414 | 0.59766 | 0.0019461 | 0.32561 | -0.013521 | -2.2623 | 4.2444e-07 | 0.61118 | 0.1295 | 1.63 | 0.031144 | 0.03063 | 0.00011 | 0.35913 | 1.9024 | 1.6691 | 0.00051982 |  | 2.9 | 0.00090317 | -0.13 | 0.030562 | 1.9625e-05 | 19.6246
3100 | 37758.7 | 28804.7 |  |  | -3.139 | 3.139 | 0.47438 | 0.03824 | -1.2182 | -0.0094712 | 0.0094712 | 0.0014313 | 0.30173 | 5.3329e-07 | -3.1295 | 0.74418 | 0.003172 | 0.42625 | -0.01352 | -1.8168 | 2.0741e-07 | 0.7577 | 0.1375 | 1.63 | 0.03329 | 0.032732 | 9.5e-05 | 0.29024 | 1.8961 | 1.6556 | 0.00055117 |  | 3.2 | 0.0010653 | -0.11 | 0.032671 | 2.276e-05 | 22.7603
3200 | 66563.4 | 28804.7 |  |  | -3.6094 | 3.6094 | 0.54547 | 0.036751 | -1.0182 | -0.0094699 | 0.0094699 | 0.0014311 | 0.26237 | 4.8885e-07 | -3.6 | 0.85132 | 0.0029152 | 0.34243 | -0.01352 | -1.5882 | 3.2257e-07 | 0.86484 | 0.1422 | 1.63 | 0.034161 | 0.033631 | 9.6e-05 | 0.28545 | 1.7502 | 1.6548 | 0.00056531 |  | 2.9 | 0.00099068 | -0.1 | 0.033574 | 2.5316e-05 | 25.3161
3300 | 95368.1 | 28804.7 |  |  | -4.0329 | 4.0329 | 0.60946 | 0.041553 | -1.0304 | -0.009469 | 0.009469 | 0.001431 | 0.2348 | 7.4787e-07 | -4.0234 | 0.95236 | 0.0035652 | 0.37435 | -0.013521 | -1.4197 | 5.4884e-07 | 0.96588 | 0.1455 | 1.63 | 0.03493 | 0.034394 | 8.9e-05 | 0.25877 | 1.7128 | 1.6504 | 0.00057648 |  | 2.9 | 0.001013 | -0.09 | 0.034341 | 2.7652e-05 | 27.6522
3400 | 124172.8 | 28804.7 |  |  | -4.4235 | 4.4235 | 0.66849 | 0.044489 | -1.0057 | -0.0094674 | 0.0094674 | 0.0014308 | 0.21403 | 7.2127e-07 | -4.414 | 1.0553 | 0.0060864 | 0.57675 | -0.013522 | -1.2813 | 6.7262e-07 | 1.0688 | 0.1478 | 1.63 | 0.035789 | 0.03534 | 9e-05 | 0.25467 | 1.4254 | 1.6498 | 0.00059044 |  | 3 | 0.0010737 | -0.09 | 0.035286 | 2.9865e-05 | 29.8646
3500 | 152977.5 | 28804.7 |  |  | -4.6946 | 4.6946 | 0.70947 | 0.048847 | -1.0405 | -0.0094675 | 0.0094675 | 0.0014308 | 0.20167 | 5.498e-07 | -4.6851 | 1.1374 | 0.0059386 | 0.52214 | -0.013522 | -1.1889 | 5.7128e-07 | 1.1509 | 0.1481 | 1.63 | 0.03638 | 0.035982 | 9.2e-05 | 0.25568 | 1.2437 | 1.6499 | 0.00060024 |  | 3 | 0.0010914 | -0.08 | 0.035933 | 3.1635e-05 | 31.635
3600 | 181782.2 | 28804.7 |  |  | -4.7776 | 4.7776 | 0.72201 | 0.049975 | -1.046 | -0.009467 | 0.009467 | 0.0014307 | 0.19815 | 3.9679e-07 | -4.7681 | 1.1872 | 0.0058361 | 0.49158 | -0.013523 | -1.1391 | 6.5061e-07 | 1.2007 | 0.1447 | 1.63 | 0.036439 | 0.03585 | 0.0001 | 0.27894 | 1.7565 | 1.6537 | 0.0006026 |  | 3 | 0.0010932 | -0.065 | 0.03581 | 3.2952e-05 | 32.9518
3700 | 210586.9 | 28804.7 |  |  | -4.9434 | 4.9434 | 0.74707 | 0.053122 | -1.0746 | -0.009466 | 0.009466 | 0.0014305 | 0.19149 | 4.5979e-07 | -4.9339 | 1.4272 | 0.0090548 | 0.63443 | -0.013523 | -0.94753 | 5.7304e-07 | 1.4407 | 0.1449 | 1.33 | 0.042312 | 0.04117 | 0.00026 | 0.63153 | 3.0366 | 1.4723 | 0.00062297 |  | 3.2 | 0.001354 | -0.15 | 0.041065 | 3.4051e-05 | 34.0505
3800 | 239391.6 | 28804.7 |  |  | -5.4795 | 5.4795 | 0.82808 | 0.057237 | -1.0446 | -0.0094657 | 0.0094657 | 0.0014305 | 0.17275 | 6.3601e-07 | -5.47 | 2.1375 | 0.014545 | 0.68048 | -0.013523 | -0.63267 | 3.9225e-07 | 2.151 | 0.1567 | 0.79 | 0.061621 | 0.0604 | 0.00042 | 0.69536 | 2.3691 | 1.0524 | 0.00064852 |  | 3 | 0.0018486 | -0.2 | 0.060195 | 3.4908e-05 | 34.9076
3900 | 268196.3 | 28804.7 |  |  | -6.0409 | 6.0409 | 0.91292 | 0.068257 | -1.1299 | -0.0094648 | 0.0094648 | 0.0014304 | 0.15668 | 5.8565e-07 | -6.0314 | 2.8938 | 0.01298 | 0.44855 | -0.013523 | -0.46732 | 7.0415e-07 | 2.9073 | 0.1691 | 0.65 | 0.081511 | 0.08051 | 0.00036 | 0.44715 | 1.5022 | 0.78895 | 0.00064308 |  | 2.8 | 0.0022823 | -0.15 | 0.080305 | 3.5668e-05 | 35.6676
4000 | 297001 | #REF! |  |  | -6.6171 | 6.6171 | 1 | 0.070879 | -1.0712 | -0.0094643 | 0.0094643 | 0.0014303 | 0.14303 | 3.3147e-07 | -6.6076 | 3.5475 | 0.017099 | 0.48202 | -0.013523 | -0.38119 | 5.7967e-07 | 3.561 | 0.1798 | 0.49 | 0.096899 | 0.09599 | 0.00032 | 0.33337 | 1.1442 | 0.59265 | 0.00057427 |  | 2.6 | 0.0025194 | -0.115 | 0.095802 | 3.675e-05 | 36.7497
λ [Å] | Steps | Step interval | Environmental conditions |  | STD - ORW-005.RA | [10 empty] | UUT - PtSi | [6 empty] | STD responsivity [A/W] | STD Uc [%] | UUT responsivity [A/W] | PtSi cert response [A/W] | PtSi cert Uc [A/W] | PtSi cert Uc [%] | Deviation between PtSi cal cert and now [%] | Combined Uc |  | Remarks
|  |  | Temp | Humidity | Light |  |  |  |  | Dark | [5 empty] | Light |  |  | Dark
[5 empty] | Average [V] |  |  | Std Dev [V] | Std Dev [%] | Average [V] |  |  |  | Std Dev [V] | Net [V] | Average [V] | Std Dev [V] | Std Dev [%] | Average [V] |  | Std Dev [V] | Net [V]
Filter no. 1: Blank LDLS |  |  | 0.2 | 1.6
2000 | -279093 | 28804.7 | 1 |  | 0.23136 |  |  | 0.00050567 | 0.21856 | 0.0094873 |  |  | 4.1007 | 1.4687e-06 | 0.22187 | -0.039875 | 0.00017453 | -0.43768 | 0.013413 | -33.6375 | 6.9582e-07 | -0.053289 | 0.1134 | 6.12 | 0.027236 | 0.02676 | 0.00029 | 1.0837 | 1.78 | 6.2152 | 0.0016928 | 18/03/2019
2100 | -250288.3 | 28804.7 | 23.5 | 54.8 | 0.66439 |  |  | 0.0041148 | 0.61933 | 0.0094771 |  |  | 1.4264 | 1.8406e-06 | 0.65492 | -0.1343 | 0.0011114 | -0.82755 | 0.013415 | -9.989 | 1.6025e-06 | -0.14771 | 0.1163 | 4.25 | 0.026231 | 0.0262 | 0.0002 | 0.76336 | 0.11796 | 4.318 | 0.0011327
2200 | -221483.6 | 28804.7 |  |  | 1.5263 |  |  | 0.016643 | 1.0904 | 0.0094703 |  |  | 0.62045 | 2.6754e-06 | 1.5169 | -0.31426 | 0.002319 | -0.73791 | 0.013415 | -4.2689 | 1.5922e-06 | -0.32768 | 0.1209 | 2.29 | 0.026117 | 0.02564 | 0.00017 | 0.66303 | 1.8598 | 2.3841 | 0.00062264
2300 | -192678.9 | 28804.7 |  |  | 3.1198 |  |  | 0.033958 | 1.0885 | 0.0094653 |  |  | 0.30339 | 1.9459e-06 | 3.1104 | -0.64288 | 0.0045682 | -0.71059 | 0.013416 | -2.0869 | 1.8241e-06 | -0.6563 | 0.127 | 2.02 | 0.026797 | 0.026 | 0.00016 | 0.61538 | 3.0669 | 2.1117 | 0.00056587
2400 | -163874.2 | 28804.7 |  |  | 4.1405 |  |  | 0.046567 | 1.1247 | 0.0094612 |  |  | 0.2285 | 1.2281e-06 | 4.1311 | -0.85327 | 0.006007 | -0.704 | 0.013416 | -1.5723 | 1.4171e-06 | -0.86669 | 0.1304 | 2 | 0.027357 | 0.02677 | 0.00017 | 0.63504 | 2.1945 | 2.0984 | 0.00057407
2500 | -135069.5 | 28804.7 |  |  | 4.8551 |  |  | 0.057181 | 1.1778 | 0.0094581 |  |  | 0.19481 | 2.2417e-06 | 4.8456 | -1.0012 | 0.0071798 | -0.71712 | 0.013415 | -1.3398 | 1.0477e-06 | -1.0146 | 0.1219 | 1.64 | 0.025524 | 0.02473 | 0.00015 | 0.60655 | 3.2123 | 1.7486 | 0.00044631
2600 | -106264.8 | 28804.7 |  |  | 5.3006 |  |  | 0.068721 | 1.2965 | 0.0094557 |  |  | 0.17839 | 1.9234e-06 | 5.2912 | -1.0869 | 0.012722 | -1.1704 | 0.013415 | -1.2342 | 2.878e-06 | -1.1004 | 0.1103 | 1.66 | 0.022938 | 0.022396 | 9.8e-05 | 0.43758 | 2.4202 | 1.7167 | 0.00039378
2700 | -77460.1 | 28804.7 |  |  | 5.9585 |  |  | 0.070069 | 1.176 | 0.0094537 |  |  | 0.15866 | 2.6459e-06 | 5.949 | -1.218 | 0.013737 | -1.1278 | 0.013415 | -1.1013 | 1.6119e-06 | -1.2315 | 0.1001 | 1.64 | 0.020721 | 0.02025 | 0.0001 | 0.49383 | 2.3258 | 1.7127 | 0.0003549
2800 | -48655.4 | 28804.7 |  |  | 7.098 |  |  | 0.0824 | 1.1609 | 0.0094535 |  |  | 0.13319 | 1.4821e-06 | 7.0885 | -1.5389 | 0.018072 | -1.1744 | 0.013413 | -0.87163 | 3.0679e-06 | -1.5523 | 0.1025 | 1.63 | 0.022446 | 0.02146 | 0.00013 | 0.60578 | 4.5942 | 1.7389 | 0.00039032
2900 | -19850.7 | 28804.7 |  |  | 0.097344 |  |  | 0.00017896 | 0.18385 | 0.0094422 |  |  | 9.6998 | 7.4449e-07 | 0.087902 | -0.011313 | 8.1043e-05 | -0.71635 | 0.0094772 | -83.7705 | 8.3535e-07 | -0.02079 | 0.1147 | 1.64 | 0.027129 | 0.02618 | 0.00013 | 0.49656 | 3.6239 | 1.7135 | 0.00046486
3000 | 8954 | 28804.7 |  |  | 0.11243 |  |  | 0.00080854 | 0.71914 | 0.0094414 |  |  | 8.3974 | 5.033e-07 | 0.10299 | -0.015589 | 0.00010012 | -0.64223 | 0.0094767 | -60.7913 | 6.6196e-07 | -0.025066 | 0.1295 | 1.63 | 0.031517 | 0.03063 | 0.00011 | 0.35913 | 2.8972 | 1.6691 | 0.00052605
3100 | 37758.7 | 28804.7 |  |  | 0.11925 |  |  | 0.00085798 | 0.71949 | 0.0094417 |  |  | 7.9177 | 9.9524e-07 | 0.10981 | -0.017303 | 0.00010012 | -0.57864 | 0.0094769 | -54.7708 | 2.0363e-06 | -0.02678 | 0.1375 | 1.63 | 0.033534 | 0.032732 | 9.5e-05 | 0.29024 | 2.449 | 1.6556 | 0.0005552
3200 | 66563.4 | 28804.7 |  |  | 0.12185 |  |  | 0.00089382 | 0.73354 | 0.0094415 |  |  | 7.7484 | 6.4526e-07 | 0.11241 | -0.017761 | 0.00013606 | -0.76606 | 0.0094766 | -53.3567 | 2.2458e-06 | -0.027238 | 0.1422 | 1.63 | 0.034456 | 0.033631 | 9.6e-05 | 0.28545 | 2.4533 | 1.6548 | 0.00057018
3300 | 95368.1 | 28804.7 |  |  | 0.12449 |  |  | 0.0009024 | 0.7249 | 0.0094406 |  |  | 7.5837 | 8.904e-07 | 0.11505 | -0.018208 | 0.00013094 | -0.71912 | 0.0094772 | -52.05 | 2.1083e-06 | -0.027685 | 0.1455 | 1.63 | 0.035014 | 0.034394 | 8.9e-05 | 0.25877 | 1.8018 | 1.6504 | 0.00057787
3400 | 124172.8 | 28804.7 |  |  | 0.12533 |  |  | 0.00093844 | 0.74879 | 0.0094399 |  |  | 7.5323 | 1.6009e-06 | 0.11589 | -0.018776 | 0.00013858 | -0.73805 | 0.0094768 | -50.4727 | 1.2859e-06 | -0.028253 | 0.1478 | 1.63 | 0.036033 | 0.03534 | 9e-05 | 0.25467 | 1.9619 | 1.6498 | 0.00059447
3500 | 152977.5 | 28804.7 |  |  | 0.12303 |  |  | 0.00090974 | 0.73943 | 0.0094392 |  |  | 7.6721 | 2.7644e-06 | 0.11359 | -0.018519 | 0.00010928 | -0.5901 | 0.0094771 | -51.1766 | 6.1238e-07 | -0.027996 | 0.1481 | 1.63 | 0.0365 | 0.035982 | 9.2e-05 | 0.25568 | 1.4393 | 1.6499 | 0.00060222
3600 | 181782.2 | 28804.7 |  |  | 0.11723 |  |  | 0.00086723 | 0.73974 | 0.0094402 |  |  | 8.0524 | 1.0148e-06 | 0.10779 | -0.017911 | 0.00011235 | -0.62725 | 0.0094763 | -52.9067 | 1.3809e-06 | -0.027388 | 0.1447 | 1.63 | 0.036765 | 0.03585 | 0.0001 | 0.27894 | 2.5516 | 1.6537 | 0.00060798
3700 | 210586.9 | 28804.7 |  |  | 0.11493 |  |  | 0.00088684 | 0.77163 | 0.0094409 |  |  | 8.2144 | 8.66e-07 | 0.10549 | -0.022677 | 0.00015259 | -0.6729 | 0.0094773 | -41.7933 | 2.9885e-06 | -0.032154 | 0.1449 | 1.33 | 0.044166 | 0.04117 | 0.00026 | 0.63153 | 7.2778 | 1.4723 | 0.00065027
3800 | 239391.6 | 28804.7 |  |  | 0.12112 |  |  | 0.00092592 | 0.76446 | 0.0094411 |  |  | 7.7947 | 1.0573e-06 | 0.11168 | -0.036418 | 0.0002536 | -0.69637 | 0.0094775 | -26.0245 | 2.9177e-06 | -0.045895 | 0.1567 | 0.79 | 0.064396 | 0.0604 | 0.00042 | 0.69536 | 6.6159 | 1.0524 | 0.00067773
3900 | 268196.3 | 28804.7 |  |  | 0.12618 |  |  | 0.00097661 | 0.77397 | 0.0094412 |  |  | 7.4821 | 8.093e-07 | 0.11674 | -0.047999 | 0.00033087 | -0.68932 | 0.0094785 | -19.7472 | 2.7066e-06 | -0.057478 | 0.1691 | 0.65 | 0.083256 | 0.08051 | 0.00036 | 0.44715 | 3.4111 | 0.78895 | 0.00065685
4000 | 297001 | #REF! |  |  | 0.12705 |  |  | 0.0009337 | 0.7349 | 0.0094412 |  |  | 7.431 | 8.4998e-07 | 0.11761 | -0.054427 | 0.00033536 | -0.61616 | 0.0094778 | -17.4138 | 2.7295e-06 | -0.063904 | 0.1798 | 0.49 | 0.097696 | 0.09599 | 0.00032 | 0.33337 | 1.7773 | 0.59265 | 0.000579
[10 empty] | y = 288.047x-855187 | [5 empty] | Step calculator | Wavelength [A] | Step count
[10 empty] | 288.047 |  |  |  |  | 968150.51 |  | 9050 | 1751638.35
855187
[8 empty] | 2000 | -279093 | 5500 |  |  |  | 729071.5 | 288047
|  |  |  | 6970 |  |  |  | 2500 | -135069.5 | 5000 |  |  |  | 585048
[10 empty] | 4500 |  |  |  | 441024.5 | 288047
[10 empty] | 3500 |  |  |  | 152977.5
|  |  |  | 9658 |  |  |  | -144023.5 |  | 9618 |  |  |  | 1915249.046
[10 empty] | 9698 |  |  |  | 1938292.806

## S_N ratio
St | 6.6171
Sd | 0.0006
σ St | 2e-05
σ Sd | 2e-05
S/N | 233927.251
Uc [%] | 0.00042748
𝑆/𝑁= (𝑆^𝑡−𝑆^𝑑 )/√(𝜎^2 (𝑆^𝑡)+𝜎^2 (𝑆^𝑑))
the higher the S/N ratio, the smaller the Uc
𝑈_𝑐=1 % 𝑓𝑜𝑟 𝑆/𝑁=100
the higher the S/N ratio, the smaller the Uc

## Sheet1
STD - ORW-005.RA
| Xenon arc source | [5 empty] | LDLS | [5 empty] | % diff
|  |  | norm | [5 empty] | norm
| light | dark | light | dark | light | dark | light | dark | light | dark | light | dark
200 | 0.0099425 | 0.0095037 | 0.0014008 | 0.0013389 | 0.0014008 | 0.0013389 | 0.23136 | 0.0094873 | 0.032595 | 0.0013366 | 0.032595 | 0.0013366 | 3.1194 | 95.7025
210 | 0.01036 | 0.0094971 | 0.0014595 | 0.001338 | 0.0014595 | 0.001338 | 0.66439 | 0.0094771 | 0.093603 | 0.0013352 | 0.093603 | 0.0013352 | 9.2144 | 98.4407
220 | 0.012528 | 0.0094922 | 0.001765 | 0.0013373 | 0.001765 | 0.0013373 | 1.5263 | 0.0094703 | 0.21504 | 0.0013342 | 0.21504 | 0.0013342 | 21.3274 | 99.1792
230 | 0.031503 | 0.0094883 | 0.0044383 | 0.0013368 | 0.0044383 | 0.0013368 | 3.1198 | 0.0094653 | 0.43954 | 0.0013335 | 0.43954 | 0.0013335 | 43.51 | 98.9902
240 | 0.11302 | 0.0094843 | 0.015922 | 0.0013362 | 0.015922 | 0.0013362 | 4.1405 | 0.0094612 | 0.58334 | 0.0013329 | 0.58334 | 0.0013329 | 56.7416 | 97.2705
250 | 0.27913 | 0.0094818 | 0.039326 | 0.0013358 | 0.039326 | 0.0013358 | 4.8551 | 0.0094581 | 0.68401 | 0.0013325 | 0.68401 | 0.0013325 | 64.468 | 94.2507
260 | 0.5296 | 0.0094791 | 0.074613 | 0.0013355 | 0.074613 | 0.0013355 | 5.3006 | 0.0094557 | 0.74678 | 0.0013322 | 0.74678 | 0.0013322 | 67.2165 | 90.0087
270 | 0.86818 | 0.0094776 | 0.12231 | 0.0013352 | 0.12231 | 0.0013352 | 5.9585 | 0.0094537 | 0.83946 | 0.0013319 | 0.83946 | 0.0013319 | 71.7144 | 85.4295
280 | 1.2794 | 0.0094757 | 0.18025 | 0.001335 | 0.18025 | 0.001335 | 7.098 | 0.0094535 | 1 | 0.0013319 | 1 | 0.0013319 | 81.9754 | 81.9754
290 | 1.862 | 0.0094746 |  |  | 0.26232 | 0.0013348 | 0.097344 | 0.0094422 |  |  | 0.013714 | 0.0013303 | 0
300 | 2.5509 | 0.0094731 |  |  | 0.35938 | 0.0013346 | 0.11243 | 0.0094414 |  |  | 0.01584 | 0.0013302 | 0
310 | 3.139 | 0.0094712 |  |  | 0.44224 | 0.0013343 | 0.11925 | 0.0094417 |  |  | 0.0168 | 0.0013302 | 0
320 | 3.6094 | 0.0094699 |  |  | 0.50851 | 0.0013342 | 0.12185 | 0.0094415 |  |  | 0.017167 | 0.0013302 | 0
330 | 4.0329 | 0.009469 |  |  | 0.56817 | 0.001334 | 0.12449 | 0.0094406 |  |  | 0.017538 | 0.00133 | 0
340 | 4.4235 | 0.0094674 |  |  | 0.6232 | 0.0013338 | 0.12533 | 0.0094399 |  |  | 0.017657 | 0.0013299 | 0
350 | 4.6946 | 0.0094675 |  |  | 0.6614 | 0.0013338 | 0.12303 | 0.0094392 |  |  | 0.017333 | 0.0013298 | 0
360 | 4.7776 | 0.009467 |  |  | 0.67309 | 0.0013338 | 0.11723 | 0.0094402 |  |  | 0.016516 | 0.00133 | 0
370 | 4.9434 | 0.009466 |  |  | 0.69645 | 0.0013336 | 0.11493 | 0.0094409 |  |  | 0.016192 | 0.0013301 | 0
380 | 5.4795 | 0.0094657 |  |  | 0.77197 | 0.0013336 | 0.12112 | 0.0094411 |  |  | 0.017064 | 0.0013301 | 0
390 | 6.0409 | 0.0094648 |  |  | 0.85106 | 0.0013334 | 0.12618 | 0.0094412 |  |  | 0.017777 | 0.0013301 | 0
400 | 6.6171 | 0.0094643 |  |  | 0.93224 | 0.0013334 | 0.12705 | 0.0094412 |  |  | 0.0179 | 0.0013301 | 0
LDLS
Output signal with LDLS and shutter closed
Xenon arc source
Output signal with Xenon-arc source and shutter closed

## Uncertainty calcCOPYMSc
Uncertainty contributions | [17 empty] | Calibration of spectral power responsivity of STD detector | [7 empty] | Drift in STD, from paper | Drift in STD | Polarization, non-uniform beam profile and repeatability | Polarization, non-uniform beam profile and repeatability | Beam divergence & vignetting divergence, negligible if 2,4° | Linearity of the standard detector | Spectral wavelength calibration | Lamp stability / Drift during cal. |  |  |  |  | Temperature dependence of STD | DVM uncertainty |  |  |  |  | Spatial uniformity of UUT | Amplifier uncertainty same gain setting used for STD & UUT | Scattered light (internal stray light) | External stray light (dark reading subtracted, use as indication of signal-noise ratio) | Repeatability in STD |  |  |  | Effect of resolution of the multimeter on the output signal of the standard detector |  |  |  | Bandwidth effects on UUT | Spatial uniformity of UUT |  |  |  | Temperature dependence of UUT |  | DVM uncertainty |  |  |  |  | Repeatability in UUT |  |  |  | Resolution of UUT
Remarks | [29 empty] | CCPR-K2.c 2003 | [8 empty] | Uncertainty |  | Accuracy: ± (0,002 % of reading + 0,0006 % of range) | [21 empty] | Uncertainty |  | Accuracy: ± (0,002 % of reading + 0,0006 % of range)
Estimated U | [17 empty] | 0,49 to 6,12 % | [9 empty] | 0.041% | 0.01% |  |  | 0,3 nm | 0.217% |  |  |  |  | See below | Range [V] | Uc [V] | Range [V] | % of reading | % of range |  | 0% | 0.0015% |  | See below |  |  |  | See below | [8 empty] | See below |  | Range [V] | Uc [V] | Range [V] | % of reading | % of range | See below |  |  |  | See below
Probability distr. | [17 empty] | Normal | [7 empty] | Normal | Normal | Normal | Normal | Normal | Normal | Normal | Rectangular |  |  |  |  | Normal | 0.0001 | 1.2e-06 | 0.1 | 0.003% | 0.003% | Normal | Normal | normal | normal | Normal |  |  |  | Rectangular |  |  |  | Rectangular | Normal |  |  |  | Rectangular |  | 0.0001 | 1.2e-06 | 0.1 | 0.003% | 0.003% | Normal |  |  |  | Rectangular
k = | [17 empty] | 2 | [7 empty] | 3.18 | 1 | 1 | 1 | 1 | 1 | 2 | 1.7321 |  |  |  |  | 1 | 0.001 | 1.2e-06 | 1 | 0.002% | 0.0006% | 1 | 1 | 1 | 1 | 1 |  |  |  | 1.7321 |  |  |  | 1.7321 | 1 |  |  |  | 1.7321 |  | 0.001 | 1.2e-06 | 1 | 0.002% | 0.0006% | 1 |  |  |  | 1.7321
Standard U | [17 empty] | See below | [7 empty] | See below | See below | See below | See below | See below | See below | See below | See below | [5 empty] | 0.01 | 1.2e-06 | 10 | 0.0015% | 0.0004% | See below | See below | See below | See below | See below |  |  |  | See below |  |  |  | See below | See below | [5 empty] | 0.01 | 1.2e-06 | 10 | 0.0015% | 0.0004% | See below |  |  |  | See below
Max std Uc [%] | [17 empty] | 0.00072416 | [14 empty] | 0.21696 | [5 empty] | 0.1 | 1.2e-06 | divisor |  | 1.7321 | 0.086 | 0 | 0.153% | 95.5867 | 1.2182 |  |  |  | 2.7865e-07 | [10 empty] | 0.1 | 1.2e-06 | divisor |  | 1.7321 | 0.77294 |  |  |  | 2.7355e-06
[39 empty] | 1 | 6e-06 |  |  | See below | [19 empty] | 1 | 6e-06 |  |  | See below
[39 empty] | 10 | 4e-05 |  |  | 0.13674 | [19 empty] | 10 | 4e-05 |  |  | 0.65522
[39 empty] | divisor | 2 | [22 empty] | divisor | 2
| Standard - Si: ORW-005.RA | [7 empty] | UUT - PtSi: ORW-523.RA | [7 empty] | Combined standard uncertainty contribution from the standard detector and reference equipment | [37 empty] | UUT Uncertainty contributions | [17 empty] | Uncertainty and remarks
Wavelength | Spectral power responsivity | Temp coeff interpolated | Corrected Spectral power responsivity | Deviation | Signal of STD (Light minus dark) | Standard Deviation in signal (Light) | S/N |  | Spectral power responsivity | PtSi spectral power responsivity | Spectral power responsivity calibration values from PTB of PtSi, corrected for deviation from calibration temperature (25,2 °C), using temperature coefficient. ΔT = 25,2-23,5 | Signal of UUT (Light minus dark) | Standard Deviation in signal (Light) | S/N |  |  | Absolute spectral power responsivity of the standard detector | [6 empty] | Drift in the standard detector |  |  |  | Polarization and non-uniform beam profile | Inter-reflections between detector and output aperture of the monochromator | Beam divergence and vignetting | Linearity of the standard detector | Spectral wavelength calibration | Source stability |  | Temperature coefficient of the standard detector |  | Temperature dependence of the standard detector |  | Effect of the multimeter on the standard detector's output signal |  |  |  |  | Spatial uniformity of the standard detector | Effect of the amplifier | Internal stray light | External stray light | Repeatability of the standard detector | Effect of resolution of the multimeter on the output signal of the standard detector |  |  |  | Combined standard uncertainty contribution BMC (k = 1) | Expanded relative uncertainty BMC (k = 2) | Bandwidth effects on UUT |  | Spatial uniformity of UUT | Temperature coefficient of UUT |  | Temperature dependence of UUT |  | Linearity of UUT | DVM uncertainty - UUT |  |  |  |  | Repeatability uncertainty contribution | Resolution of UUT |  |  |  | Combined uncertainty standard contribution BMC (k = 1) | Relative Expanded Uncertainty BMC (k = 2) | Combined uncertainty contribution UUT included (k = 1) | Expanded Uncertainty UUT included (k = 2) |  | Deviation between PTB calibration & NMISA calculated spectral power responsivity | Verification statement | Remarks
[8 empty] | CIE 202:2011 page 7 | [6 empty] | % noise of signal, insignificant if less than 1/3 of uncertainty | CIE 202:2011 page 7 | Spectral power responsivity Certificate Uncertainty (k=2) |  | Spectral power responsivity Certificate Uncertainty | Spectral power responsivity certificate standard uncertainty | Sensitivity Coefficient | Standard uncertainty contribution | Standard uncertainty contribution | Drift in response (per annum) | Standard deviation of drift rate σ(λ) (Prediction contribution U p.a.) | for k = 3,18 | [7 empty] | ? | -2.5 | [13 empty] | Resolution of STD | Semi range of resolution of STD | Resolution of STD standard uncertainty contribution |  |  |  | Standard uncertainty contribution |  | Standard uncertainty contribution | 1.7 |  | Standard uncertainty contribution |  | Standard uncertainty contribution | Standard uncertainty |  | Standard uncertainty contribution for accuracy |  | Combined standard uncertainty contribution for DVM |  | Resolution of UUT | Semi range of resolution of UUT | Resolution of UUT standard uncertainty contribution | Resolution of STD standard uncertainty contribution
Source/ reference | NPL Certificate 2017060356 | Interpolated from CCPR-k2.b 2004 report | ΔT = 2 °C |  | Measured | Calculated | Kostkowski | 100 S/N ratio = 1 % Uc | Calculated | PTB Certificate 73328 18 PTB |  | Measured | Calculated | Kostkowski | Kostkowski page 276 | 100 S/N ratio = 1 % Uc | NPL Certificate 2017060356 | [6 empty] | From paper: Predicting wavlength dependent responsivity-drift is silicon, Littler et al (NewRad proceedings) |  |  |  | CCPR-K2.c 2003 worst case from participants Uaver p50 | CCPR-K2.c 2003 | PTB Uncertainty values | CCPR-K2.c 2003 | Certificate OR\SR-5491 & calculated | Measured |  | CCPR-K2.c 2003, Si detector manufacturer specifications & estimate |  | Optical Radiometry book & estimate |  | k=2 | k=1 | k=1,732 | k=1 | DVM certificate and accuracy specification | CCPR-K2.c 2003 | Common factor for both STD & UUT, therefore no contribution | McPherson specification | Calculated from dark reading of std | Caclulated from measured data | Caclulated from measured data | [5 empty] | CCPR-K2.c, detector of same type |  | CCPR-K2.c 2003 | CCPR-K2.c 2003 |  | CCPR-K2.c 2003 |  | CCPR-K2.c 2003 | k=2 | k=1 | k=1,732 | k=1 | DVM certificate and accuracy specification | Caclulated from measured data | Caclulated from measured data | [7 empty] | rounded up
Probability ditribution | [16 empty] | Gaussian
Coverage factor, k | [16 empty] | k = 2 | k = 1
| sSTD_ [A/W] |  |  |  | VSTD_ [V] | [V] |  | rel Uc | sUUT_ [A/W] | [A/W] | [A/W] | VUUT_ [V] | [V] |  |  | rel Uc | us [%] | [6 empty] | [%] | [%] | [%] | uSd [%] | upnb [%] | uRefl [%] | ubdv [%] | uSlin [%] | uλ [%] | uss [%] |  | ΔT [°C] |  | U(Td) [%] | uStempd [%] | U(DVM_u) [%] | umc [%] | U(DVM_acc) [%] | uma [%] | uSm [%] | uSsu [%] | ua [%] | uisl [%] | uesl [%] | uSr [%] | [V] | [V] | [V] | uSres [%] | UBMC [%] |  | U(bw) [%] | U(bw) [%] | U(SU) [%] | ΔT = 25,2-23,5 [°C] [%/deg C] |  | U(T) [%] | U(T) [%] | U(Lin) [%] | U(DVM_u) [%] | U(DVM_u) [%] | U(DVM_acc) [%] | U(DVM_acc) [%] | U(DVM) [%] | [%] | [V] | [V] | [V] | U(ResUUT) [%] | U [%] | U [%] | U [%] | U [%] |  | [%]
[nm] |  | [1/K] | [A/W] | [%] |  |  | [1] | [%] | [5 empty] | [1] |  | [%] |  |  | U(sSTD) [A/W] | U(sSTD) [A/W] | c_sUUT,sSTD [1] | Uc(sSTD) [A/W] | Uc(sSTD) [%] | [11 empty] | [%/deg C] | uStempc [%] | [44 empty] | Resolution |  | En
200 | 0.1134 |  |  |  | 0.0004388 | 1.1497e-06 | 275.7348 | 0.36267 | 0.023665 | 0.02676 | 0.026719 | 9.1572e-05 | 4.6328e-07 | 117.0282 | 0.85449 | 0.85449 | 6.12 | 3.06 | 0.0069401 | 0.00347 | 0.20869 | 0.00072416 | 0.63859 | -0.01076 | 0.05206 | 0.16555 | 0.05316 | 0.041 | 0.01 | 0.005 | 0.02 | 0.061189 | 0.21696 | 494.4341 | 0.05 | -0.125 | 0.02 | 0.02 | 0.27347 | 0.13674 | 3.0132e-06 | 1.7397e-06 | 0.13674 | 0.086 | 0 | 0.00153 | 95.5867 | 0.011563 | 1e-11 | 5e-12 | 2.8868e-12 | 2.9034e-08 | 95.6361 | 191.2723 | 0.029 | 0.016743 | 0.124% | -0.09 | -0.153 | 0.016% | 0.009238% | 0.02 | 1.3104 | 0.65522 | 3.0027e-06 | 1.7336e-06 | 0.65522 | -0.0034518 | 1e-10 | 5e-11 | 2.8868e-11 | -2.1508e-07 | 95.6361 | 191.2723 | 95.6385 | 191.277 | 191.3 | 11.4287 | OK | This range is in noise level, will not use these values | 0.045272 | 191.3 | 0.06745
210 | 0.1163 |  |  |  | 0.00086255 | 2.1869e-06 | 381.6072 | 0.26205 | 0.024629 | 0.0262 | 0.026155 | 0.00018267 | 5.3452e-07 | 239.1074 | 0.41822 | 0.41822 | 4.25 | 2.125 | 0.0049428 | 0.0024714 | 0.21177 | 0.00052337 | 0.45002 | -0.012099 | 0.050125 | 0.1594 | 0.051564 | 0.041 | 0.01 | 0.005 | 0.02 | 0.058521 | 0.21696 | 251.5288 | 0.05 | -0.125 | 0.02 | 0.02 | 0.13912 | 0.069561 | 3.0259e-06 | 1.747e-06 | 0.069561 | 0.086 | 0 | 0.00153 | 91.674 | 0.02111 | 1e-10 | 5e-11 | 2.8868e-11 | 2.7865e-07 | 91.699 | 183.3981 | 0.015 | 0.0086603 | 0.124% | -0.1 | -0.17 | 0.017% | 0.009815% | 0.02 | 0.65694 | 0.32847 | 3.0055e-06 | 1.7352e-06 | 0.32847 | -0.0040096 | 1e-10 | 5e-11 | 2.8868e-11 | -2.1654e-07 | 91.699 | 183.3981 | 91.6998 | 183.3996 | 183.4 | 5.835 | OK |  | 0.04517 | 183.4 | 0.033787
220 | 0.1209 |  |  |  | 0.003036 | 1.5399e-05 | 196.9234 | 0.50781 | 0.025554 | 0.02564 | 0.025596 | 0.0006417 | 1.1351e-06 | 507.9181 | 0.19688 | 0.19688 | 2.29 | 1.145 | 0.0027686 | 0.0013843 | 0.21137 | 0.00029259 | 0.24201 | -0.013697 | 0.048228 | 0.15336 | 0.050135 | 0.041 | 0.01 | 0.005 | 0.02 | 0.050304 | 0.21696 | 71.4616 | 0.05 | -0.125 | 0.02 | 0.02 | 0.039526 | 0.019763 | 3.0911e-06 | 1.7846e-06 | 0.019763 | 0.086 | 0 | 0.00153 | 75.7667 | 0.12292 | 1e-10 | 5e-11 | 2.8868e-11 | 2.3042e-07 | 75.776 | 151.5519 | 0.01 | 0.0057735 | 0.124% | -0.1 | -0.17 | 0.016% | 0.009238% | 0.02 | 0.187 | 0.093501 | 3.0193e-06 | 1.7432e-06 | 0.093501 | -0.0088176 | 1e-10 | 5e-11 | 2.8868e-11 | -2.2424e-07 | 75.776 | 151.5519 | 75.7762 | 151.5524 | 151.6 | 0.16551 | OK |  | 0.03874 | 151.6 | 0.0010935
230 | 0.127 |  |  |  | 0.022015 | 0.00014225 | 154.7615 | 0.64616 | 0.026466 | 0.026 | 0.025956 | 0.0045878 | 1.5588e-05 | 294.1921 | 0.33991 | 0.33991 | 2.02 | 1.01 | 0.0025654 | 0.0012827 | 0.20839 | 0.0002673 | 0.21048 | -0.015623 | 0.046368 | 0.14745 | 0.048929 | 0.041 | 0.01 | 0.005 | 0.02 | 0.063507 | 0.21696 | 9.8549 | 0.05 | -0.125 | 0.02 | 0.02 | 0.0054508 | 0.0027254 | 3.6605e-06 | 2.1134e-06 | 0.0027254 | 0.086 | 0 | 0.00153 | 30.1183 | 0.45154 | 1e-10 | 5e-11 | 2.8868e-11 | 9.1633e-08 | 30.1399 | 60.2798 | 0.005 | 0.0028868 | 0.124% | -0.1 | -0.17 | 0.017% | 0.009815% | 0.02 | 0.026157 | 0.013078 | 3.1376e-06 | 1.8115e-06 | 0.013078 | -0.17459 | 1e-11 | 5e-12 | 2.8868e-12 | -3.2333e-08 | 30.1399 | 60.2798 | 30.1409 | 60.2818 | 60.3 | 1.9647 | OK |  | 0.015959 | 60.3 | 0.031953
240 | 0.1304 |  |  |  | 0.10353 | 0.00073212 | 141.4131 | 0.70715 | 0.027128 | 0.02677 | 0.026724 | 0.021538 | 6.1994e-05 | 347.407 | 0.28785 | 0.28785 | 2 | 1 | 0.002608 | 0.001304 | 0.20803 | 0.00027128 | 0.20803 | -0.017969 | 0.044546 | 0.14166 | 0.048034 | 0.041 | 0.01 | 0.005 | 0.02 | 0.034274 | 0.21696 | 2.0956 | 0.05 | -0.125 | 0.02 | 0.02 | 0.0011591 | 0.00057954 | 6.1059e-06 | 3.5253e-06 | 0.00057955 | 0.086 | 0 | 0.00153 | 8.3921 | 0.6478 | 1e-09 | 5e-10 | 2.8868e-10 | 2.5543e-07 | 8.4808 | 16.9615 | 0.019 | 0.01097 | 0.124% | -0.1 | -0.17 | 0.017% | 0.009815% | 0.02 | 0.0055716 | 0.0027858 | 3.6461e-06 | 2.1051e-06 | 0.0027858 | 0.77294 | 1e-11 | 5e-12 | 2.8868e-12 | 3.5992e-08 | 8.4808 | 16.9615 | 8.5176 | 17.0353 | 17.1 | 1.508 | OK |  | 0.0046388 | 17.1 | 0.08682
250 | 0.1219 |  |  |  | 0.26965 | 0.0016792 | 160.5841 | 0.62273 | 0.025184 | 0.02473 | 0.024692 | 0.055708 | 9.2275e-05 | 603.7025 | 0.16564 | 0.16564 | 1.64 | 0.82 | 0.0019992 | 0.00099958 | 0.20659 | 0.00020651 | 0.16941 | -0.02086 | 0.042763 | 0.13598 | 0.047579 | 0.041 | 0.01 | 0.005 | 0.02 | 0.14818 | 0.21696 | 0.80458 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00044502 | 0.00022251 | 1.109e-05 | 6.4026e-06 | 0.0002226 | 0.086 | 0 | 0.00153 | 3.3969 | 0.60157 | 1e-09 | 5e-10 | 2.8868e-10 | 1.0342e-07 | 3.5576 | 7.1153 | 0.003 | 0.0017321 | 0.124% | -0.09 | -0.153 | 0.015% | 0.00866% | 0.02 | 0.0021541 | 0.001077 | 4.6712e-06 | 2.6969e-06 | 0.001077 | 0.21872 | 1e-10 | 5e-11 | 2.8868e-11 | 6.8423e-08 | 3.5576 | 7.1153 | 3.5677 | 7.1354 | 7.2 | 1.9901 | OK |  | 0.0018132 | 7.2 | 0.27009
260 | 0.1103 |  |  |  | 0.52013 | 0.002895 | 179.6616 | 0.5566 | 0.022741 | 0.022396 | 0.022364 | 0.10724 | 0.00038182 | 280.8568 | 0.35605 | 0.35605 | 1.66 | 0.83 | 0.001831 | 0.00091549 | 0.20617 | 0.00018875 | 0.17113 | -0.024469 | 0.041016 | 0.13043 | 0.04776 | 0.041 | 0.01 | 0.005 | 0.02 | 0.17063 | 0.21696 | 0.41712 | 0.02 | -0.05 | 0.02 | 0.02 | 0.0011536 | 0.00057678 | 1.6403e-05 | 9.47e-06 | 0.00057686 | 0.086 | 0 | 0.00153 | 1.7899 | 0.54664 | 1e-09 | 5e-10 | 2.8868e-10 | 5.4508e-08 | 2.0694 | 4.1387 | 0.009 | 0.0051962 | 0.124% | -0.085 | -0.1445 | 0.014% | 0.008083% | 0.02 | 0.001119 | 0.00055951 | 6.2171e-06 | 3.5894e-06 | 0.00055952 | 0.40742 | 1e-10 | 5e-11 | 2.8868e-11 | 3.0803e-08 | 2.0694 | 4.1387 | 2.1141 | 4.2283 | 4.3 | 1.6877 | OK | Ask for accreditation in this range | 0.00097787 | 4.3 | 0.38406
270 | 0.1001 |  |  |  | 0.8587 | 0.0046691 | 183.9103 | 0.54374 | 0.02064 | 0.02025 | 0.020222 | 0.17706 | 0.000586 | 302.1564 | 0.33095 | 0.33095 | 1.64 | 0.82 | 0.0016416 | 0.00082082 | 0.2062 | 0.00016925 | 0.16908 | -0.029035 | 0.039308 | 0.125 | 0.048869 | 0.041 | 0.01 | 0.005 | 0.02 | 0.057581 | 0.21696 | 0.25266 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00069873 | 0.00034936 | 2.3174e-05 | 1.338e-05 | 0.00034962 | 0.086 | 0 | 0.00153 | 1.0917 | 0.53781 | 1e-09 | 5e-10 | 2.8868e-10 | 3.3251e-08 | 1.4895 | 2.979 | 0.015 | 0.0086603 | 0.124% | -0.08 | -0.136 | 0.012% | 0.006928% | 0.02 | 0.00067772 | 0.00033886 | 8.3119e-06 | 4.7989e-06 | 0.0003389 | 0.35832 | 1e-09 | 5e-10 | 2.8868e-10 | 1.7651e-07 | 1.4895 | 2.979 | 1.5382 | 3.0764 | 3.1 | 2.0672 | OK |  | 0.00063986 | 3.1 | 0.64549
280 | 0.1025 |  |  |  | 1.2699 | 0.013756 | 92.319 | 1.0832 | 0.021977 | 0.02146 | 0.02142 | 0.27228 | 0.0014241 | 191.1934 | 0.52303 | 0.52303 | 1.63 | 0.815 | 0.0016707 | 0.00083537 | 0.21441 | 0.00017911 | 0.17474 | -0.034897 | 0.037638 | 0.11969 | 0.051326 | 0.041 | 0.01 | 0.005 | 0.02 | 0.2365 | 0.21696 | 0.17084 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00047247 | 0.00023624 | 3.1398e-05 | 1.8128e-05 | 0.00023693 | 0.086 | 0 | 0.00153 | 0.74064 | 1.0752 | 1e-08 | 5e-09 | 2.8868e-09 | 2.2564e-07 | 1.577 | 3.154 | 0.014 | 0.0080829 | 0.124% | -0.11 | -0.187 | 0.019% | 0.01097% | 0.02 | 0.00044072 | 0.00022036 | 1.1168e-05 | 6.4481e-06 | 0.00022046 | 0.55036 | 1e-09 | 5e-10 | 2.8868e-10 | 1.1156e-07 | 1.577 | 3.154 | 1.6808 | 3.3617 | 3.4 | 2.6005 | OK |  | 0.00074721 | 3.4 | 0.73443
290 | 0.1147 |  |  |  | 1.8525 | 0.018634 | 99.4169 | 1.0059 | 0.026723 | 0.02618 | 0.026122 | 0.4316 | 0.001489 | 289.8619 | 0.34499 | 0.34499 | 1.64 | 0.82 | 0.0018811 | 0.00094054 | 0.23298 | 0.00021913 | 0.19105 | -0.042529 | 0.036005 | 0.1145 | 0.055723 | 0.041 | 0.01 | 0.005 | 0.02 | 0.29071 | 0.21696 | 0.11712 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00032389 | 0.00016194 | 4.305e-05 | 2.4855e-05 | 0.00016384 | 0.086 | 0 | 0.00153 | 0.50885 | 1.0007 | 1e-08 | 5e-09 | 2.8868e-09 | 1.5504e-07 | 1.4422 | 2.8845 | 0.021 | 0.012124 | 0.124% | -0.13 | -0.221 | 0.022% | 0.0127% | 0.02 | 0.00027803 | 0.00013902 | 1.5948e-05 | 9.2076e-06 | 0.00013932 | 0.35615 | 1e-09 | 5e-10 | 2.8868e-10 | 6.9047e-08 | 1.4422 | 2.8845 | 1.5021 | 3.0042 | 3.1 | 2.3015 | OK |  | 0.00082842 | 3 | 0.71694
300 | 0.1295 | -9.3e-05 | 0.12948 | 0.018603 | 2.5414 | 0.02538 | 100.1342 | 0.99866 | 0.031144 | 0.03063 | 0.030562 | 0.61118 | 0.0019461 | 314.063 | 0.31841 | 0.31841 | 1.63 | 0.815 | 0.0021109 | 0.0010554 | 0.24049 | 0.00025382 | 0.196 | -0.052594 | 0.03441 | 0.10942 | 0.06285 | 0.041 | 0.01 | 0.005 | 0.02 | 0.15971 | 0.21696 | 0.085369 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00023609 | 0.00011805 | 5.6828e-05 | 3.2809e-05 | 0.00012252 | 0.086 | 0 | 0.00153 | 0.37137 | 0.99495 | 1e-08 | 5e-09 | 2.8868e-09 | 1.1317e-07 | 1.3715 | 2.7431 | 0.002 | 0.0011547 | 0.124% | -0.13 | -0.221 | 0.021% | 0.01212% | 0.02 | 0.0009817 | 0.00049085 | 1.8224e-05 | 1.0521e-05 | 0.00049096 | 0.32561 | 1e-09 | 5e-10 | 2.8868e-10 | 4.8301e-08 | 1.3715 | 2.7431 | 1.427 | 2.854 | 2.9 | 1.9024 | OK |  | 0.00090317 | 2.9 | 0.63904
310 | 0.1375 | -5.9765e-05 | 0.13748 | 0.011954 | 3.1295 | 0.03824 | 81.8391 | 1.2219 | 0.03329 | 0.032732 | 0.032671 | 0.7577 | 0.003172 | 238.8678 | 0.41864 | 0.41864 | 1.63 | 0.815 | 0.0022412 | 0.0011206 | 0.24211 | 0.00027132 | 0.19732 | -0.065976 | 0.032853 | 0.10447 | 0.073703 | 0.041 | 0.01 | 0.005 | 0.02 | 0.056636 | 0.21696 | 0.069325 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00019172 | 9.5861e-05 | 6.8591e-05 | 3.9601e-05 | 0.00010372 | 0.086 | 0 | 0.00153 | 0.30173 | 1.2182 | 1e-08 | 5e-09 | 2.8868e-09 | 9.1964e-08 | 1.5191 | 3.0381 | 0.001 | 0.00057735 | 0.124% | -0.11 | -0.187 | 0.018% | 0.01039% | 0.02 | 0.00079187 | 0.00039594 | 2.1154e-05 | 1.2213e-05 | 0.00039613 | 0.42625 | 1e-09 | 5e-10 | 2.8868e-10 | 3.8791e-08 | 1.5191 | 3.0381 | 1.5889 | 3.1778 | 3.2 | 1.8961 | OK |  | 0.0010653 | 3.2 | 0.57921
320 | 0.1422 | -2.2e-05 | 0.14219 | 0.0044002 | 3.6 | 0.036751 | 97.9563 | 1.0209 | 0.034161 | 0.033631 | 0.033574 | 0.86484 | 0.0029152 | 296.6646 | 0.33708 | 0.33708 | 1.63 | 0.815 | 0.0023179 | 0.0011589 | 0.24024 | 0.00027842 | 0.19579 | -0.083731 | 0.031334 | 0.099641 | 0.089402 | 0.041 | 0.01 | 0.005 | 0.02 | 0.031185 | 0.21696 | 0.060266 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00016667 | 8.3334e-05 | 7.7999e-05 | 4.5033e-05 | 9.4724e-05 | 0.086 | 0 | 0.00153 | 0.26237 | 1.0182 | 1e-08 | 5e-09 | 2.8868e-09 | 7.9978e-08 | 1.3558 | 2.7117 | 0.001 | 0.00057735 | 0.124% | -0.1 | -0.17 | 0.016% | 0.009238% | 0.02 | 0.00069377 | 0.00034689 | 2.3297e-05 | 1.345e-05 | 0.00034715 | 0.34243 | 1e-09 | 5e-10 | 2.8868e-10 | 3.3909e-08 | 1.3558 | 2.7117 | 1.4089 | 2.8177 | 2.9 | 1.7502 | OK |  | 0.00099068 | 2.8 | 0.59036
330 | 0.1455 | 2.0795e-05 | 0.14551 | 0.0041587 | 4.0234 | 0.041553 | 96.8245 | 1.0328 | 0.03493 | 0.034394 | 0.034341 | 0.96588 | 0.0035652 | 270.9195 | 0.36911 | 0.36911 | 1.63 | 0.815 | 0.0023716 | 0.0011858 | 0.24007 | 0.00028468 | 0.19565 | -0.1067 | 0.029852 | 0.09493 | 0.1108 | 0.041 | 0.01 | 0.005 | 0.02 | 0.035248 | 0.21696 | 0.053924 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00014913 | 7.4564e-05 | 8.6468e-05 | 4.9922e-05 | 8.9733e-05 | 0.086 | 0 | 0.00153 | 0.2348 | 1.0304 | 1e-08 | 5e-09 | 2.8868e-09 | 7.1581e-08 | 1.3617 | 2.7233 | 0.001 | 0.00057735 | 0.124% | -0.09 | -0.153 | 0.015% | 0.00866% | 0.02 | 0.0006212 | 0.0003106 | 2.5318e-05 | 1.4617e-05 | 0.00031094 | 0.37435 | 1e-09 | 5e-10 | 2.8868e-10 | 3.0312e-08 | 1.3617 | 2.7233 | 1.4206 | 2.8412 | 2.9 | 1.7128 | OK |  | 0.001013 | 2.9 | 0.57845
340 | 0.1478 | 5.3e-05 | 0.14782 | 0.010599 | 4.414 | 0.044489 | 99.2158 | 1.0079 | 0.035789 | 0.03534 | 0.035286 | 1.0688 | 0.0060864 | 175.6074 | 0.56945 | 0.56945 | 1.63 | 0.815 | 0.0024091 | 0.0012046 | 0.24214 | 0.00029168 | 0.19735 | -0.13435 | 0.028408 | 0.090339 | 0.13732 | 0.041 | 0.01 | 0.005 | 0.02 | 0.037258 | 0.21696 | 0.049152 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00013593 | 6.7966e-05 | 9.428e-05 | 5.4432e-05 | 8.7076e-05 | 0.086 | 0 | 0.00153 | 0.21403 | 1.0057 | 1e-08 | 5e-09 | 2.8868e-09 | 6.526e-08 | 1.3422 | 2.6843 | 0.002 | 0.0011547 | 0.124% | -0.09 | -0.153 | 0.015% | 0.00866% | 0.02 | 0.00056137 | 0.00028068 | 2.7376e-05 | 1.5806e-05 | 0.00028113 | 0.57675 | 1e-07 | 5e-08 | 2.8868e-08 | 2.7355e-06 | 1.3422 | 2.6843 | 1.469 | 2.9379 | 3 | 1.4254 | OK |  | 0.0010737 | 3 | 0.46682
350 | 0.1481 | 5.8586e-05 | 0.14812 | 0.011716 | 4.6851 | 0.048847 | 95.9137 | 1.0426 | 0.03638 | 0.035982 | 0.035933 | 1.1509 | 0.0059386 | 193.797 | 0.516 | 0.516 | 1.63 | 0.815 | 0.002414 | 0.001207 | 0.24564 | 0.0002965 | 0.2002 | -0.16239 | 0.027003 | 0.085868 | 0.16462 | 0.041 | 0.01 | 0.005 | 0.02 | 0.016219 | 0.21696 | 0.046307 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00085376 | 0.00042688 | 0.00011028 | 6.3668e-05 | 0.0004316 | 0.086 | 0 | 0.00153 | 0.20167 | 1.0405 | 1e-08 | 5e-09 | 2.8868e-09 | 6.1491e-08 | 1.3691 | 2.7382 | 0.005 | 0.0028868 | 0.124% | -0.08 | -0.136 | 0.013% | 0.007506% | 0.02 | 0.00052134 | 0.00026067 | 2.9018e-05 | 1.6753e-05 | 0.00026121 | 0.52214 | 1e-08 | 5e-09 | 2.8868e-09 | 2.5381e-07 | 1.3691 | 2.7382 | 1.4717 | 2.9435 | 3 | 1.2437 | OK |  | 0.0010914 | 3 | 0.40804
360 | 0.1447 | 3.6e-05 | 0.14471 | 0.0071995 | 4.7681 | 0.049975 | 95.4099 | 1.0481 | 0.036439 | 0.03585 | 0.03581 | 1.2007 | 0.0058361 | 205.7428 | 0.48604 | 0.48604 | 1.63 | 0.815 | 0.0023586 | 0.0011793 | 0.25183 | 0.00029698 | 0.20524 | -0.18132 | 0.025634 | 0.081517 | 0.18312 | 0.041 | 0.01 | 0.005 | 0.02 | 0.090397 | 0.21696 | 0.045501 | 0.02 | -0.05 | 0.02 | 0.02 | 0.0008389 | 0.00041945 | 0.00011152 | 6.4387e-05 | 0.00042436 | 0.086 | 0 | 0.00153 | 0.19815 | 1.046 | 1e-08 | 5e-09 | 2.8868e-09 | 6.0423e-08 | 1.378 | 2.7561 | 0.013 | 0.0075056 | 0.124% | -0.065 | -0.1105 | 0.01% | 0.005774% | 0.02 | 0.00049969 | 0.00024984 | 3.0015e-05 | 1.7329e-05 | 0.00025045 | 0.49158 | 1e-08 | 5e-09 | 2.8868e-09 | 2.4315e-07 | 1.378 | 2.7561 | 1.4674 | 2.9348 | 3 | 1.7565 | OK |  | 0.0010932 | 2.9 | 0.57301
370 | 0.1449 | -1.1015e-05 | 0.1449 | 0.0022031 | 4.9339 | 0.053122 | 92.8796 | 1.0767 | 0.042312 | 0.04117 | 0.041065 | 1.4407 | 0.0090548 | 159.1143 | 0.62848 | 0.62848 | 1.33 | 0.665 | 0.0019272 | 0.00096358 | 0.29201 | 0.00028137 | 0.19419 | -0.18132 | 0.024304 | 0.077287 | 0.18294 | 0.041 | 0.01 | 0.005 | 0.02 | 0.46638 | 0.21696 | 0.043972 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00081071 | 0.00040536 | 0.00011401 | 6.5823e-05 | 0.00041067 | 0.086 | 0 | 0.00153 | 0.19149 | 1.0746 | 1e-08 | 5e-09 | 2.8868e-09 | 5.8396e-08 | 1.3943 | 2.7887 | 0.013 | 0.0075056 | 0.124% | -0.15 | -0.255 | 0.024% | 0.01386% | 0.02 | 0.00041645 | 0.00020823 | 3.4815e-05 | 2.01e-05 | 0.00020919 | 0.63443 | 1e-08 | 5e-09 | 2.8868e-09 | 2.0226e-07 | 1.3943 | 2.7887 | 1.5531 | 3.1063 | 3.2 | 3.0366 | OK |  | 0.001354 | 3.1 | 0.90445
380 | 0.1567 | -7.2e-05 | 0.15668 | 0.014402 | 5.47 | 0.057237 | 95.5676 | 1.0464 | 0.061621 | 0.0604 | 0.060195 | 2.151 | 0.014545 | 147.884 | 0.67621 | 0.67621 | 0.79 | 0.395 | 0.0012379 | 0.00061897 | 0.39324 | 0.0002434 | 0.15533 | -0.16239 | 0.023012 | 0.073177 | 0.16401 | 0.041 | 0.01 | 0.005 | 0.02 | 0.52948 | 0.21696 | 0.039663 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00073126 | 0.00036563 | 0.00012205 | 7.0466e-05 | 0.00037236 | 0.086 | 0 | 0.00153 | 0.17275 | 1.0446 | 1e-08 | 5e-09 | 2.8868e-09 | 5.2683e-08 | 1.2821 | 2.5642 | 0.004 | 0.0023094 | 0.124% | -0.2 | -0.34 | 0.032% | 0.01848% | 0.02 | 0.00027894 | 0.00013947 | 4.9021e-05 | 2.8302e-05 | 0.00014231 | 0.68048 | 1e-08 | 5e-09 | 2.8868e-09 | 1.3505e-07 | 1.2821 | 2.5642 | 1.4909 | 2.9819 | 3 | 2.3691 | OK |  | 0.0018486 | 3 | 0.75225
390 | 0.1691 | -0.0001339 | 0.16905 | 0.026787 | 6.0314 | 0.068257 | 88.3631 | 1.1317 | 0.081511 | 0.08051 | 0.080305 | 2.9073 | 0.01298 | 223.9843 | 0.44646 | 0.44646 | 0.65 | 0.325 | 0.0010991 | 0.00054957 | 0.48203 | 0.00026491 | 0.15666 | -0.13435 | 0.021757 | 0.069187 | 0.1361 | 0.041 | 0.01 | 0.005 | 0.02 | 0.32073 | 0.21696 | 0.035971 | 0.02 | -0.05 | 0.02 | 0.02 | 0.0006632 | 0.0003316 | 0.00013047 | 7.5327e-05 | 0.00034005 | 0.086 | 0 | 0.00153 | 0.15668 | 1.1299 | 1e-08 | 5e-09 | 2.8868e-09 | 4.7787e-08 | 1.2601 | 2.5202 | 0.002 | 0.0011547 | 0.124% | -0.15 | -0.255 | 0.025% | 0.01443% | 0.02 | 0.00020638 | 0.00010319 | 6.4146e-05 | 3.7035e-05 | 0.00010963 | 0.44855 | 1e-08 | 5e-09 | 2.8868e-09 | 9.9757e-08 | 1.2601 | 2.5202 | 1.3618 | 2.7236 | 2.8 | 1.5022 | OK |  | 0.0022823 | 2.7 | 0.52211
400 | 0.1798 | -0.00018 | 0.17974 | 0.036013 | 6.6076 | 0.070879 | 93.2238 | 1.0727 | 0.096899 | 0.09599 | 0.095802 | 3.561 | 0.017099 | 208.2516 | 0.48019 | 0.48019 | 0.49 | 0.245 | 0.00088102 | 0.00044051 | 0.53892 | 0.0002374 | 0.13204 | -0.1067 | 0.02054 | 0.065317 | 0.10866 | 0.041 | 0.01 | 0.005 | 0.02 | 0.22242 | 0.21696 | 0.032834 | 0.02 | -0.05 | 0.02 | 0.02 | 0.00060536 | 0.00030268 | 0.00013911 | 8.0317e-05 | 0.00031316 | 0.086 | 0 | 0.00153 | 0.14303 | 1.0712 | 1e-08 | 5e-09 | 2.8868e-09 | 4.3626e-08 | 1.1613 | 2.3227 | 0.001 | 0.00057735 | 0.124% | -0.115 | -0.1955 | 0.019% | 0.01097% | 0.02 | 0.00016849 | 8.4246e-05 | 7.722e-05 | 4.4583e-05 | 9.5316e-05 | 0.48202 | 1e-08 | 5e-09 | 2.8868e-09 | 8.1375e-08 | 1.1613 | 2.3227 | 1.2727 | 2.5453 | 2.6 | 1.1442 | OK |  | 0.0025194 | 2.6 | 0.43165
[45 empty] | with filter | 4e-10
[45 empty] | without filter | 4.8e-05
0.00083333
Note on CG20: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: En values will be very small, because Uc is very large

## MScTables
| Absolute spectral power responsivity of the standard detector | Drift in the standard detector | Spatial uniformity of the standard detector | Linearity of the standard detector | Temperature coefficient of the standard detector | Temperature dependence of the standard detector | Effect of the amplifier | Effect of the multimeter on the standard detector's output signal |  |  |  | Effect of resolution of the multimeter on the output signal of the standard detector | Source stability | Polarisation and non-uniform beam profile | Inter-reflections between detector and output aperture of the monochromator | Beam divergence and vignetting | Internal stray light | Effect of external stray light on the standard detector | Spectral wavelength calibration | Repeatability of the standard detector | Bandwidth effects on the test detector |  | Spatial uniformity of the test detector | Linearity of the test detector | Temperature coefficient of the test detector | Temperature dependence of the test detector | Effect of the multimeter on the test detector's output signal |  |  | Effect of resolution of the multimeter on the output signal of the test detector | Repeatability of the test detector |  | Combined standard uncertainty | Expanded Uncertainty UUT (k = 2) | Wavelength | Combined standard uncertainty contribution, BMC (k = 1) | Expanded relative uncertainty, BMC (k = 2) | Combined standard uncertainty contribution including the test detector (k = 1) | Expanded relative uncertainty including the test detector (k = 2) | Combined uncertainty standard contribution BMC (k = 1) | Relative Expanded Uncertainty BMC (k = 2) | Combined uncertainty contribution UUT included (k = 1) | Expanded Uncertainty UUT included (k = 2) |  | Deviation between PTB calibration & NMISA calculated spectral power responsivity | Verification statement | Remarks
Wavelength | us | uSd | uSsu | uSlin | uStempc | uStempd | ua | umc | uma | uSm | Wavelength | uSres | uss | upnb | urefl | ubdv | uisl | uSesl | uλ | uSr | uTb | Wavelength | uTsu | uTlin | uTtempc | uTtempd | uTmc | uTma | uTm | uTres | uTr | Wavelength | uc(y) | U |  | UBMC |  | U |  | U [%] | U [%] | U [%] | U [%] |  | [%]
[nm] | [%] | [%] | [%] | [%] | [%] | [%] | [%] | [%] | [%] | [%] | [nm] | [%] | [%] | [%] | [%] | [%] | [%] | [%] | [%] | [%] | [%] | [nm] | [%] | [%] | [%] | [%] | [%] | [%] | [%] | [%] | [%] | [nm] | [%] | [%] | [nm] | [%] |  | [%]
260 | 0.83 | 0.04776 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 0.00057678 | 9.47e-06 | 0.00057686 | 260 | 5.4508e-08 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 1.7899 | 0.17063 | 0.54664 | 0.0051962 | 260 | 0.124 | 0.02 | 0.1445 | 0.0080829 | 0.00055951 | 3.5894e-06 | 0.00055952 | 3.0803e-08 | 0.40742 | 260 | 2.1141 | 4.3 | 260 | 450.3416 | 900.6833 | 450.3416 | 900.7 | 2.0694 | 4.1387 | 2.1141 | 4.2283 | 4.3 | 1.6877
270 | 0.82 | 0.048869 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 0.00034936 | 1.338e-05 | 0.00034962 | 270 | 3.3251e-08 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 1.0917 | 0.057581 | 0.53781 | 0.0086603 | 270 | 0.124 | 0.02 | 0.136 | 0.0069282 | 0.00033886 | 4.7989e-06 | 0.0003389 | 1.7651e-07 | 0.35832 | 270 | 1.5382 | 3.1 | 270 | 467.6574 | 935.3149 | 467.6574 | 935.4 | 1.4895 | 2.979 | 1.5382 | 3.0764 | 3.1 | 2.0672
280 | 0.815 | 0.051326 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 0.00023624 | 1.8128e-05 | 0.00023693 | 280 | 2.2564e-07 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 0.74064 | 0.2365 | 1.0752 | 0.0080829 | 280 | 0.124 | 0.02 | 0.187 | 0.01097 | 0.00022036 | 6.4481e-06 | 0.00022046 | 1.1156e-07 | 0.55036 | 280 | 1.6808 | 3.4 | 280 | 484.9775 | 969.955 | 484.9775 | 970 | 1.577 | 3.154 | 1.6808 | 3.3617 | 3.4 | 2.6005
290 | 0.82 | 0.055723 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 0.00016194 | 2.4855e-05 | 0.00016384 | 290 | 1.5504e-07 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 0.50885 | 0.29071 | 1.0007 | 0.012124 | 290 | 0.124 | 0.02 | 0.221 | 0.012702 | 0.00013902 | 9.2076e-06 | 0.00013932 | 6.9047e-08 | 0.35615 | 290 | 1.5021 | 3.1 | 290 | 502.2972 | 1004.5944 | 502.2972 | 1004.6 | 1.4422 | 2.8845 | 1.5021 | 3.0042 | 3.1 | 2.3015
300 | 0.815 | 0.06285 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 0.00011805 | 3.2809e-05 | 0.00012252 | 300 | 1.1317e-07 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 0.37137 | 0.15971 | 0.99495 | 0.0011547 | 300 | 0.124 | 0.02 | 0.221 | 0.012124 | 0.00049085 | 1.0521e-05 | 0.00049096 | 4.8301e-08 | 0.32561 | 300 | 1.427 | 2.9 | 300 | 519.6173 | 1039.2345 | 519.6173 | 1039.3 | 1.3715 | 2.7431 | 1.427 | 2.854 | 2.9 | 1.9024
310 | 0.815 | 0.073703 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 9.5861e-05 | 3.9601e-05 | 0.00010372 | 310 | 9.1964e-08 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 0.30173 | 0.056636 | 1.2182 | 0.00057735 | 310 | 0.124 | 0.02 | 0.187 | 0.010392 | 0.00039594 | 1.2213e-05 | 0.00039613 | 3.8791e-08 | 0.42625 | 310 | 1.5889 | 3.2 | 310 | 536.938 | 1073.8761 | 536.938 | 1073.9 | 1.5191 | 3.0381 | 1.5889 | 3.1778 | 3.2 | 1.8961
320 | 0.815 | 0.089402 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 8.3334e-05 | 4.5033e-05 | 9.4724e-05 | 320 | 7.9978e-08 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 0.26237 | 0.031185 | 1.0182 | 0.00057735 | 320 | 0.124 | 0.02 | 0.17 | 0.0092376 | 0.00034689 | 1.345e-05 | 0.00034715 | 3.3909e-08 | 0.34243 | 320 | 1.4089 | 2.9 | 320 | 554.258 | 1108.5161 | 554.258 | 1108.6 | 1.3558 | 2.7117 | 1.4089 | 2.8177 | 2.9 | 1.7502
330 | 0.815 | 0.1108 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 7.4564e-05 | 4.9922e-05 | 8.9733e-05 | 330 | 7.1581e-08 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 0.2348 | 0.035248 | 1.0304 | 0.00057735 | 330 | 0.124 | 0.02 | 0.153 | 0.0086603 | 0.0003106 | 1.4617e-05 | 0.00031094 | 3.0312e-08 | 0.37435 | 330 | 1.4206 | 2.9 | 330 | 571.5785 | 1143.157 | 571.5785 | 1143.2 | 1.3617 | 2.7233 | 1.4206 | 2.8412 | 2.9 | 1.7128
340 | 0.815 | 0.13732 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 6.7966e-05 | 5.4432e-05 | 8.7076e-05 | 340 | 6.526e-08 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 0.21403 | 0.037258 | 1.0057 | 0.0011547 | 340 | 0.124 | 0.02 | 0.153 | 0.0086603 | 0.00028068 | 1.5806e-05 | 0.00028113 | 2.7355e-06 | 0.57675 | 340 | 1.469 | 3 | 340 | 588.8989 | 1177.7978 | 588.8989 | 1177.8 | 1.3422 | 2.6843 | 1.469 | 2.9379 | 3 | 1.4254
350 | 0.815 | 0.16462 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 0.00042688 | 6.3668e-05 | 0.0004316 | 350 | 6.1491e-08 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 0.20167 | 0.016219 | 1.0405 | 0.0028868 | 350 | 0.124 | 0.02 | 0.136 | 0.0075056 | 0.00026067 | 1.6753e-05 | 0.00026121 | 2.5381e-07 | 0.52214 | 350 | 1.4717 | 3 | 350 | 606.2194 | 1212.4388 | 606.2194 | 1212.5 | 1.3691 | 2.7382 | 1.4717 | 2.9435 | 3 | 1.2437
360 | 0.815 | 0.18312 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 0.00041945 | 6.4387e-05 | 0.00042436 | 360 | 6.0423e-08 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 0.19815 | 0.090397 | 1.046 | 0.0075056 | 360 | 0.124 | 0.02 | 0.1105 | 0.0057735 | 0.00024984 | 1.7329e-05 | 0.00025045 | 2.4315e-07 | 0.49158 | 360 | 1.4674 | 3 | 360 | 623.5399 | 1247.0798 | 623.5399 | 1247.1 | 1.378 | 2.7561 | 1.4674 | 2.9348 | 3 | 1.7565
370 | 0.665 | 0.18294 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 0.00040536 | 6.5823e-05 | 0.00041067 | 370 | 5.8396e-08 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 0.19149 | 0.46638 | 1.0746 | 0.0075056 | 370 | 0.124 | 0.02 | 0.255 | 0.013856 | 0.00020823 | 2.01e-05 | 0.00020919 | 2.0226e-07 | 0.63443 | 370 | 1.5531 | 3.2 | 370 | 640.8606 | 1281.7211 | 640.8606 | 1281.8 | 1.3943 | 2.7887 | 1.5531 | 3.1063 | 3.2 | 3.0366
380 | 0.395 | 0.16401 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 0.00036563 | 7.0466e-05 | 0.00037236 | 380 | 5.2683e-08 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 0.17275 | 0.52948 | 1.0446 | 0.0023094 | 380 | 0.124 | 0.02 | 0.34 | 0.018475 | 0.00013947 | 2.8302e-05 | 0.00014231 | 1.3505e-07 | 0.68048 | 380 | 1.4909 | 3 | 380 | 658.1808 | 1316.3617 | 658.1808 | 1316.4 | 1.2821 | 2.5642 | 1.4909 | 2.9819 | 3 | 2.3691
390 | 0.325 | 0.1361 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 0.0003316 | 7.5327e-05 | 0.00034005 | 390 | 4.7787e-08 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 0.15668 | 0.32073 | 1.1299 | 0.0011547 | 390 | 0.124 | 0.02 | 0.255 | 0.014434 | 0.00010319 | 3.7035e-05 | 0.00010963 | 9.9757e-08 | 0.44855 | 390 | 1.3618 | 2.8 | 390 | 675.5011 | 1351.0023 | 675.5011 | 1351.1 | 1.2601 | 2.5202 | 1.3618 | 2.7236 | 2.8 | 1.5022
400 | 0.245 | 0.10866 | 0.086 | 0.02 | 0.05 | 0.02 | 0 | 0.00030268 | 8.0317e-05 | 0.00031316 | 400 | 4.3626e-08 | 0.21696 | 0.041 | 0.01 | 0.005 | 0.00153 | 0.14303 | 0.22242 | 1.0712 | 0.00057735 | 400 | 0.124 | 0.02 | 0.1955 | 0.01097 | 8.4246e-05 | 4.4583e-05 | 9.5316e-05 | 8.1375e-08 | 0.48202 | 400 | 1.2727 | 2.6 | 400 | 692.8214 | 1385.6428 | 692.8214 | 1385.7 | 1.1613 | 2.3227 | 1.2727 | 2.5453 | 2.6 | 1.1442