## McPherson 2035D
UNCERTAINTY BUDGET MATRIX (UBM) | [11 empty] | Certificate No |  |  |  | OR\SR-5491 |  |  |  |  | #VALUE!
[12 empty] | Procedure No |  |  |  | SR-0007 |  |  |  |  | #VALUE!
Reference: Guide to the Expression of Uncertainty in Measurement, issued by BIPM, IEC, IFCC, ISO, IUPAC, IUPAP, OIML - ISO 1995 (ISBN 92-67-10188-9) | [17 empty] | Distribution | value |  | #VALUE!
Description: | Monochromator wavelength calibration |  | Type & Serial Number | McPherson 2035, 1 200 g/mm blazed at 300 nm |  |  |  |  | Range: | 200 nm to 400 nm | [5 empty] | Metrologist |  | - |  |  | #VALUE!
[16 empty] | I Rabe, RH Sieberhagen |  | Normal k = 1 | 1 |  | #VALUE!
[18 empty] | Normal k = 2 | 2 |  | #VALUE!
Symbol | Input Quantity | Estimated Input Quantity | Estimated Uncertainty |  | Probability Distribution (N, R, T, U) | k= | Divisor factor | Standard Uncertainty | Sensitivity Coefficient |  | Standard Uncertainty Contribution | Reliability | 1/v | Degrees of freedom calculated | Degrees of Freedom | Remarks |  | Normal k = 2,576 | 2.576 |  | #VALUE!
| (Source of Uncertainty) | [16 empty] | Normal k = 3 | 3 |  | #VALUE!
| (Xi) | (xi) |  | Unit |  | ▼ |  | U(Xi) | Ci | Unit | Ui (y) | % |  | Calculated | v |  |  | Rectangular √3 | 1.7321 |  | #VALUE!
▼ Standards and Reference Equipment (Uncorrelated) ▼ | [10 empty] | nm | [6 empty] | Rectangular √12 | 3.4641 |  | #VALUE!
λref | Theoretical value of line source in vacuum | 200 - 400 nm | 0.00017321 | nm | Normal k = 1 |  | 1 | 0.00017321 | 1 |  | 0.00017321 | 100 | 0 | 1000 | infinite | NIST ASD, see spreadsheet. Assume quoted value is standard uncertainty. |  | Triangular √6 | 2.4495 |  | #VALUE!
[18 empty] | U-Shaped √2 | 1.4142 |  | #VALUE!
n_edlen | Uncertainty in Edlen equation [n] | 1.0002 | 7.3905e-08 | 1 | Normal k = 2 |  | 2 | 3.6953e-08 | -3286.7815 | nm / 1 | -0.00012146 | 100 | 0 | 1000 | infinite | Calculated from NIST EMT and data, see sheet "Uc in Edlen eq" |  | k = ? |  |  | #VALUE!
n_temp | Lab Temperature influence on index of refraction | 23.3 | 4 | °C | Triangular √6 |  | 2.4495 | 1.633 | 0.00035275 | nm / °C | 0.00057604 | 100 | 0 | 1000 | infinite | Calculated from NIST EMT and data, see spreadsheet |  | - |  |  | #VALUE!
n_humidity | Lab Humidity influence on index of refraction | 52.5 | 30 | %RH | U-Shaped √2 |  | 1.4142 | 21.2132 | 4.5e-06 | nm / %RH | 9.5459e-05 | 100 | 0 | 1000 | infinite | Calculated from NIST EMT and data, see spreadsheet |  | Rectangular √3 | 1.7321 |  | #VALUE!
n_air pressure | Lab Air Pressure influence on index of refraction | 86.773 | 4 | kPa | Rectangular √3 |  | 1.7321 | 2.3094 | 0.0011722 | nm / kPa | 0.0027072 | 100 | 0 | 1000 | infinite | Calculated from NIST EMT and data, see spreadsheet |  | Rectangular √12 | 3.4641 |  | #VALUE!
[5 empty] | - | [15 empty] | #VALUE!
[5 empty] | - | [14 empty] | #VALUE! | #VALUE!
-
-
-
-
Res / Std | Resolution of Standard / Equipment (If applicable) |  |  |  | - | [10 empty] | Not applicable
Data / Std | Type "B" Evaluation Range of the results (Rectangular) | [14 empty] | Not applicable
| Type "A" Evaluation Exp Std Dev of the Mean (ESDM) |  |  |  | Normal K = 1 |  | 1 | [8 empty] | No of Readings
▼ Standards and Reference Equipment (Correlated) ▼ | [5 empty] | NOTE! ONLY CHANGE BLUE CELLS - All OTHER CELLS (WHITE) ARE PROTECTED
▼ Unit Under Test / Calibration (Uncorrelated) ▼ | [5 empty] | NOTE! ONLY CHANGE BLUE CELLS - All OTHER CELLS (WHITE) ARE PROTECTED
| Wavelength calibration equation offset |  | 286.5196 | steps | Normal k = 1 |  | 1 | 286.5196 | 0.00034717 | nm / steps | 0.09947 | 100 | 0 | 1000 | infinite | From verification of wavelength calibration equation
| Wavelength reproducibility |  | 0.005 | nm | Rectangular √3 |  | 1.7321 | 0.0028868 | 1 |  | 0.0028868 | 100 | 0 | 1000 | infinite | Equipment specification
| Mechanical resolution |  | 0.025 | nm | Rectangular √3 |  | 1.7321 | 0.014434 | 1 |  | 0.014434 | 100 | 0 | 1000 | infinite | Equipment specification is 0,05 nm. 2035D spec for gratings.
Res / UUT | Resolution of UUT (If applicable) |  | 0.1 | nm | Rectangular √3 |  | 1.7321 | 0.057735 | 1 | 1 | 0.057735 | 100 | 0 | 1000 | infinite | 0,2 nm used for scans
Data / UUT | Type "B" Evaluation Range of the results (Rectangular)
| Type "A" Evaluation Exp Std Dev of the Mean (ESDM) |  | 144.9876 | steps | Normal K = 1 |  | 1 | 144.9876 | 0.00034717 | nm / steps | 0.029061 | 100 | 0.5 |  | 2 | No of Readings | 3 | Type "A" Evaluation Exp Std Deviation "s" | 0.050335
▼ Unit Under Test / Calibration (Correlated) ▼ | [5 empty] | NOTE! ONLY CHANGE BLUE CELLS - All OTHER CELLS (WHITE) ARE PROTECTED | [11 empty] | Type "A" Evaluation Exp Std Dev of the Mean (ESDM) | 0.029061
-
About UBM | TOTAL COMBINED UNCERTAINTY | [9 empty] | nm | [6 empty] | Probability
Best Measurement Capability (Excluding UUT contribution) |  |  |  | Combined Uncertainty (Normal) |  |  |  | ▼ Level of Confidence ▼ |  |  | 0.0027775 | Veff |  | 1000 | infinite | Checked and Approved By: |  | 68,27 % K = 1 | 0.3173
|  |  |  | Expanded Uncertainty |  |  |  | 95,45 % K = 2 |  |  | 0.005562 | k = |  | 0.0455 | 2.0025 |  |  | 90,00 % K = 1,645 | 0.1
[18 empty] | 95,00 % K = 1,960 | 0.05
Uncertainty of Measurement (Including UUT contribution) |  |  |  | Combined Uncertainty (Normal) |  |  |  | ▼ Level of Confidence ▼ |  |  | 0.11957 | Veff |  | 573.1392 | 573.1392 |  |  | 95,45 % K = 2 | 0.0455
|  |  |  | Expanded Uncertainty |  |  |  | 95,45 % K = 2 |  |  | 0.23966 | k = |  | 0.0455 | 2.0044 |  |  | 99,00 % K = 2,576 | 0.01
Review | Review Notes | [8 empty] | was | 0.2441 | Verified |  |  |  | Date |  | 99,73 % K = 3 | 0.0027
Note on A4: Description of Unit Under Test (UUT)
Note on A6: Mathematical Formula or Model of the measurement
Note on A7: Symbols as used in the Mathematical Model
Note on A10: This section of the UBM is for entering all sources of certainty related to the Standards Should any source be correlated than these must be either uncorrelated or calculated separately
Note on A23: Resolution of the standard
Note on A24: Select Data entry either as a Type "B" (Top row) or a Type "A" (Bottom row) Evaluation. Note! Type "A" can be entered either as Experimental Standard Deviation or Estimated Standard Deviation of the Mean (ESDM) select from the drop list on the right
Note on A26: This section of the UBM is for entering all sources of certainty related to the Standards which are linked such as 2 ppm of reading + 3 ppm of range + 2 LSD
Note on A31: This section of the UBM is for entering all sources of certainty related to the UUT Should any source be correlated than these must be either uncorrelated or calculated separately
Note on A44: Select Data entry either as a Type "B" (Top row) or a Type "A" (Bottom row) Evaluation. Note! Type "A" can be entered either as Experimental Standard Deviation or Estimated Standard Deviation of the Mean (ESDM) select from the drop list on the right
Note on A46: This section of the UBM is for entering all sources of certainty related to the UUT which are linked
Note on A51: Designed by B van Oostrom © Copyright dc-Low Frequency Laboratory CSIR-NML
Note on B7: (Xi) Description of the ith input quantity on which the measurand (Y) depends Note! Xi may be a physical quantity or the random variable (see 4.1.1. note 1 of the GUM)
Note on C7: xi Estimate of the input quantity Xi Note! When xi is determined from the arithmetic mean or average of n independent repeated observations, xi =Xbari .
Note on D4: Type or Model Number of UUT
Note on D7: Estimated uncertainty of the input estimate xi
Note on D43: Irma Rabe: this should be divided by 2 for semi-range
Note on E9: Unit of the Estimated Uncertainty. Example: ppm, %, xE-X V, Ω, A, W, F, H, etc.
Note on E53: Select Confidence Level in next column on the right
Note on E56: Select Confidence Level in next column on the right
Note on F7: Probability Distribution Type "B" Evaluation Normal k = 1 Normal k = 2 Normal k = 2,576 Normal k = 3 Rectangular √3 Rectangular √12 Triangular √6 U-shape √2 k=? Enter value in column "G"
Note on G7: Enter in this column the Coverage factor if these are different from the once in the drop list of column "F"
Note on H7: Dividing Factor of the Probability Distribution selected in the column on the left. (Type "B" Evaluation)
Note on I7: u(xi) Standard uncertainty of the input estimate xi that estimates input quantity Xi , equal to the positive square root of u2(xi) Note! When xi is determined from the arithmetic mean or average of n independent repeated observations, u(xi) = s (Xbari) is a standard uncertainty obtained from a type "A" evaluation. U(xi)/PD
Note on I52: Select Confidence Level Below 68,27 % - k = 1 90,00 % - k = 1,645 95,00 % - k = 1,960 95,45 % - k = 2 99,00 % - k = 2,576 99,73 % - k = 3
Note on I55: Select Confidence LevelBelow 68,27 % - k = 1 90,00 % - k = 1,645 95,00 % - k = 1,960 95,45 % - k = 2 99,00 % - k = 2,576 99,73 % - k = 3
Note on J4: Range or Nominal Value for which this UBM is applicable
Note on J7: (Ci) Partial derivative or Sensitivity Coefficient: (Ci=∂f/∂xi)
Note on J9: Value of the Sensitivity coefficient (Ci) Partial derivative or Sensitivity Coefficient: (Ci=∂f/∂xi)
Note on K9: Unit of the Sensitivity Coefficient. µV/V µA/A µΩ/Ω mK/K etc.
Note on L7: U(xi) Actual calculated value of the associated Uncertainty U(xi) at 68% Confidence Interval. [U(xi)/PD]*Ci
Note on L10: Unit of the Standard Uncertainty Contribution
Note on L52: Best Measuring Capability (BMC) Combined (RSS) Uncertainty of all Standard Uncertainties at 68,27% Confidence Level
Note on L53: Expanded Uncertainty, U of all standards at the Level of Confidence selected in the column on the left U = k uc(y)
Note on L55: UNCERTAINTY OF MEASUREMENT Combined (RSS) Uncertainty of all sources of uncertainties at 68,27% Confidence Level
Note on L56: Expanded Uncertainty, U of allsources of uncertainty at the Level of Confidence selected in the column on the left U = k uc(y)
Note on M1: Certificate Number to which this UBM is applicable
Note on M2: Procedure which support this UBM
Note on M7: R Reliability of the Uncertainty Component (Type "B" Evaluation)
Note on P7: v Calculated Degrees of Freedom v ≈ 1/2 * [Δu(xi) / u(xi)]-2 if R = [Δu(xi) / u(xi)] * 100% v ≈ 1/2 * [100/2]2
Note on P25: Calculated Degree of Freedom for ESDM of the Type A Evaluation. (N-1)
Note on P45: Calculated Degree of Freedom for ESDM of the Type A Evaluation. (N-1)
Note on P52: V(eff) Effective Degrees of Freedom. uc4(y) Veff = --------------- n ui4 (y) Σ ------- 1 vi
Note on P53: Student "t" factor associated with the Confidence Level selected in the column on the left
Note on P55: V(eff) Effective Degrees of Freedom uc4(y) Veff = --------------- n ui4 (y) Σ ------- 1 vi
Note on P56: Student "t" factor Associated with the Confidence Level selected in the column on the left
Note on Q4: Metrologist performing the calibration and/or calculations
Note on Q7: Any remark such as where was the uncertainty from derived or obtained. Example (Certificate, Handbook, Calculated data, etc.)
Note on Q25: (n) Number of Data inputs or Measurements.
Note on Q45: (n) Number of Data inputs or Measurements.
Note on Q52: Identification of the Metrologist who has checked Uncertainty Budget Matrix for correctness

## NIST ASD Uc
Peak | True wavelength in vacuum [Å] | Uncertainty in vacuum [nm]
Refractive index @ 300 nm [1]
1 | 253.7283 | 0.0001
2 | 296.815 | 0.0001
3 | 435.956 | 0.0001
| Combined Uc [nm] | 0.00017321
Usl

## Uc in Edlen eq
Hg spectral line peak in vacuum [nm] | Corrected peak [nm] | Refractive index [n] | Corresponding monochromator steps |  | Hg spectral line peak in vacuum [nm] | Corrected peak [nm] | Refractive index [n] | Corresponding monochromator steps
302.2384 | 302.1509 | 1.0003 | 140000 |  | 2537.283 | 2536.6756 | 1.0002 | -4.301e-10
435.956 | 435.8343 | 1.0003 | 524800 |  | 2968.149 | 2967.4384 | 1.0002
546.2268 | 546.076 | 1.0003 | 842100 |  | 4359.56 | 4358.5541 | 1.0002
|  | 1.0003 |  |  |  |  | 1.0002
Hg spectral line peak in vacuum [nm] | Corrected peak [nm] | Refractive index [n] | Corresponding monochromator steps |  | Hg spectral line peak in vacuum [nm] | Corrected peak [nm] | Refractive index [n] | Corresponding monochromator steps
302.2384 | 301.8491 | 1.0013 | 140000 |  | 2537.283 | 2536.6756 | 1.0002 | 0 | 1.0002 | 2536.6755
435.956 | 435.3989 | 1.0013 | 524800 |  | 2968.149 | 2967.4384 | 1.0002 | 310710 | 1.0002 | 2967.4383
546.2268 | 545.5305 | 1.0013 | 842100 |  | 4359.56 | 4358.5541 | 1.0002 | 387819 | 1.0002 | 4358.5539
| diff in wavelength | diff in n | sens coef |  |  | diff in wavelength | diff in n | sens coef
| 0.000999 | 0.0010003 | 0.99871 |  |  | 4.3e-10 | 4.301e-10 | 0.99976 | 4.299e-08 | 4.3e-10 | 99.976
| 0.000999 | 0.0010003 | 0.99872 |  |  | 4.3e-10 | 4.301e-10 | 0.99976 | 4.299e-08 | 4.3e-10 | 99.976
| 0.000999 | 0.0010003 | 0.99873 |  |  | 4.2e-10 | 4.201e-10 | 0.99977 | 4.199e-08 | 4.2e-10 | 99.9769
ave | 0.000999 | 0.0010003 | 0.99872 |  | ave | 4.2667e-10 | 4.2677e-10 | 0.99976
[6 empty] | 1.0908e-06 | 4.301e-10 | 2536.068 | 0.00010905 | 4.3e-08 | 2536.0682
[6 empty] | 1.276e-06 | 4.301e-10 | 2966.728 | 0.00012757 | 4.3e-08 | 2966.7279
[6 empty] | 1.8306e-06 | 4.201e-10 | 4357.5493 | 0.00018302 | 4.2e-08 | 4357.5483
Peak | True wavelength in vacuum [Å] | Wavelength in lab air (24 °C, 50 %RH and 86 kPa) [Å] | Wavelength in lab air (26 °C, 65 %RH and 88 kPa) [Å] | Wavelength in lab air (22 °C, 35 %RH and 84 kPa) [Å]
[6 empty] | Hg spectral line peak in vacuum [nm] | Corrected peak [nm] | Refractive index [n] | Uncertainty in Edlen Eq. [1]
| Refractive index @ 300 nm [1] | 1.0002 | 1.0002 | 1.0002 |  | 2537.283 | 2536.6756 | 1.0002 | 4.3e-08
| Uc in refractive index [1] | 4.3e-08 | 4.4e-08 | 4.4e-08 |  | 2968.149 | 2967.4384 | 1.0002 | 4.3e-08
1 | 2537.283 | 2536.6756 | 2536.6662 | 2536.6852 |  | 4359.56 | 4358.5541 | 1.0002 | 4.2e-08
2 | 2968.149 | 2967.4384 | 2967.4274 | 2967.4497 |  |  |  | 1.0002 | 7.3905e-08
3 | 4359.56 | 4358.5541 | 4358.5386 | 4358.57 |  | Hg spectral line peak in vacuum [nm] | Corrected peak [nm] | Refractive index [n] | Corresponding monochromator steps
| Uc in refractive index [1] | 4.2e-08 |  |  |  | 2537.283 | 2536.6755 | 1.0002 | 0
[6 empty] | 2968.149 | 2967.4383 | 1.0002 | 310710
[6 empty] | 4359.56 | 4358.5539 | 1.0002 | 387819
[6 empty] | 2537.283 | 2536.6757 | 1.0002 | 0
[6 empty] | 2968.149 | 2967.4386 | 1.0002 | 310710
[6 empty] | 4359.56 | 4358.5543 | 1.0002 | 387819
[7 empty] | diff in wavelength [nm] | diff in n [1] | sens coef [nm/1]
[7 empty] | -0.00010905 | 4.3e-08 | -2536.0682
[7 empty] | -0.00012757 | 4.3e-08 | -2966.7279
[7 empty] | -0.00018302 | 4.2e-08 | -4357.5483
[6 empty] | ave | -0.00013988 | 4.2667e-08 | -3286.7815

## Uc in T
Hg spectral line peak in vacuum [nm] | at 26 degrees C | at 22 degrees C | Uc in calculation | diff in wavelength | dif in deg C | Sens Coef
435.956 | 435.8561 | 435.8547 | 4.2e-08 | 0.001411 | 4 | 0.00035275
|  |  |  | 0.001411 | 4 | 0.00035275

## Uc in %RH
Ar spectral line peak in vacuum [nm] | at 35 %RH | at 65 %RH | Uc in calculation | diff in wavelength | dif in %RH | Sens Coef [nm/%RH]
435.956 | 435.8553 | 435.8555 | 4.1e-08 | 0.000135 | 30 | 4.5e-06
|  |  |  | 0.000135 | 30 | 4.5e-06

## Uc in Pressure
Ar spectral line peak in vacuum [nm] | at 88 kPa | at 84 kPa | Uc in calculation | diff in wavelength | dif in kPa | Sens Coef
435.956 | 435.8531 | 435.8578 | 0,000000044 | 0.004689 | 4 | 0.0011722
|  |  |  | 0.004689 | 4 | 0.0011722

## ESD
at 435 nm | steps (before new origin)
scan 1 | -312254.0223 | y | -312254.0223
scan 2 | -311967.5027 | m | 288.0467
| 286.5196 | c | -855187.487
|  | x | 1884.8797
|  | y | -311967.5027
|  | m | 288.0467
|  | c | -855187.487
|  | x | 1885.8744
|  | diff in x (Å) | 0.9947
|  | diff in x (nm) | 0.09947
|  | diff in steps | 286.5196
|  | sens coef | 0.00034717
|  | y | 0
|  | m | 144.028
|  | c | -1003276
|  | x | 2968.9191
|  | y | 0
|  | m | 144.028
|  | c | -1003276
|  | x | 2968.9191
|  | diff in x (Å) | 0
|  | diff in x (nm) | 0
|  | diff in steps | 0
|  | sens coef | #DIV/0!

## Uc in step resolution
|  |  | nm | steps
y | -729611.1821 | -729561.1821 | 435.956 | -729611.1821
m | 288.047
c | -855187
x | 435.956 | 436.1296
diff in x | 0.17358
diff in steps | 50
sens coef | 0.0034717