## 5 Sep
Chopper | 3502 | 2 blade wheel
Amplifier | ORE-546.RA | 10^3 | FBW
Bewteen Femto amp & lock-in, used differential cable
LIA settings | 7214
Input limit
Sensitivity
Time const
Coupling | AC
Input shell | Float
Voltage input mode
Input connector | A-B
Input device | FET
Notch filter | 50 Hz
AC gain | OFF
Ref Source | ExtR RCU
Output slope | 12 dB Oct
Sync time const | OFF
Time const mode | Normal
DVM | ORE-111.RA
laser power controller, HeNe laser | 300 uW
DC | ORW-518.RA | [5 empty] | ORW-517.RA | [5 empty] | ORW-006.RA
| S [A/W] | 0.50782 |  |  |  |  | S [A/W] | 0.50794 | [5 empty] | S [A/W] | 0.349 | 0.35 |  |  |  | Respons from ORW-517.RA | Respons from ORW-518.RA
| light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] | Gain | Power
| ave | std dev | ave | std dev |  |  | ave | std dev | ave | std dev |  |  | ave | std dev | ave | std dev
| 0.10013 | 5.572e-06 | 0.00015753 | 1.3428e-06 | 0.099969 | 3 | 0.10014 | 4.5742e-06 | 0.00015961 | 1.1498e-06 | 0.099979 | 1 | 1.3991 | 0.00013235 | 0.00028543 | 1.2199e-05 |  | 10^5 | 65uW
[12 empty] | 2 | 0.069152 | 4.8695e-06 | 0.00015711 | 8.8121e-07 | 0.068995 | 10^3 | 300uW | 0.35053 | 0.35048
0.69016
AC (LIA); 17.5 Hz chopping freq; 100 mV sens; time const 2s | ORW-518.RA | [5 empty] | ORW-517.RA | [5 empty] | ORW-006.RA
| S [A/W] | 0.50782 |  |  |  |  | S [A/W] | 0.50794 | [5 empty] | S [A/W] | 0.349 | 0.35 |  |  |  | Respons from ORW-517.RA | Respons from ORW-518.RA
| light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] | Gain | Power
| ave | std dev | ave | std dev |  |  | ave | std dev | ave | std dev |  |  | ave | std dev | ave | std dev
| 4.3549 | 0.00033924 | 0.00023837 | 1.3734e-05 | 4.3546 |  | 4.3555 | 0.00010822 | 0.00023485 | 1.6071e-05 | 4.3552 |  | 2.7918 | 0.00047665 | 0.00022783 | 1.9169e-05 | 2.7916 | 10^3 | 300uW | 0.32558 | 0.32554 | 0.35764
| 0.043549 |  | 2.3837e-06 |  | 0.043546 |  | 0.043555 |  | 2.3485e-06 |  | 0.043552 |  | 0.027918 |  | 2.2783e-06 |  | 0.027916 |  |  |  | 0.64106 | 0.92885 | 0.90032
19.736
21.682
| 27.93 | [19 empty] | 0.91025
0.64135
RATIOS
| ORW-518.RA | ORW-517.RA | ORW-006.RA
| 0.4356 | 0.43562 | 0.4046
| 0.8712 | 0.87123 | 0.80921

## 6 Sep
Equipment settings & environmental conditions
Temperature [°C] |  | 23.8 | 23.6 |  |  |  |  | Verification:
Relative humidity [%RH] |  | 42.3 | 42.3 |  |  |  |  | Previous certificates for ORW-522.RA
Chopper |  | 3502 | 2 blade wheel | 17,5 Hz |  |  |  | 2005: OR\SR-3592 |  | 2005: OR\SR-3592 |  | 2003: OR\RA-2899
Amplifier |  | ORE-546.RA | 10^3 | FBW |  |  |  | Wavelength [nm] | Response [A/W] | Wavelength [nm] | Response [A/W] | Wavelength [nm] | Response [A/W]
Bewteen Femto amp & lock-in, used differential cable | [7 empty] | 630 | 0.334 | 630 | 0.334 | 630 | 0.347
LIA settings |  | 7214 | [5 empty] | 640 | 0.339 | 633 | 0.3355 | 632.8 | 0.34896 | 0.57811 |  | 0.44 | 0.0015443
Input limit | [7 empty] | U [%] | 2.6 | 636 | 0.337 | 636 | 0.3512 | 0.45644 |  | 0.63 | 0.0022112
Sensitivity |  | 100 mV | [7 empty] | 639 | 0.3385 | 639 | 0.3533
Time const |  | 2s | [5 empty] | 2003: OR\RA-2899 |  | U [A/W] | 0.008723 | U [A/W] | 0.009073
Coupling |  | AC | [5 empty] | 630 | 0.347 | Upper limit | 0.34422 | Upper limit | 0.35803
Input shell |  | Float | [5 empty] | 640 | 0.354 | Lower limit | 0.32678 | Lower limit | 0.33989
Input mode |  | Voltage | [5 empty] | U [%] | 2.6 | Result of measurement is within the stated uncertainty of previous certificate OR\RA-2899.
Input connector |  | A-B
Input device |  | FET
Notch filter |  | 50 Hz | [9 empty] | Dark Uc [%] |  |  |  | Repeatablitity
AC gain |  | OFF | [9 empty] | ORW-518.RA | ORW-522.RA | ORW-517.RA |  | ORW-522.RA |  | ORW-518.RA | ORW-517.RA
Ref Source |  | ExtR RCU | [9 empty] | 0.15316 | 0.21681 | 0.15055 |  | std dev | 6.7406e-06 | 0.0051897 | 0.0043346
Output slope |  | 12 dB Oct | [10 empty] | 0.21787 |  |  | esdm | 4.7663e-06
Sync time const |  | OFF | [13 empty] | ave | 0.068884
Time const mode |  | Normal | [13 empty] | % esdm | 0.0069194
DVM |  | ORE-111.RA | average of 20 readings
laser power controller, HeNe laser |  | 300 μW
DC Femto on DC setting |  | ORW-518.RA | [5 empty] | ORW-006.RA
|  | S [A/W] | 0.50782 | [5 empty] | S [A/W] | 0.349 | 0.35 |  | Respons from ORW-518.RA
|  | light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V]
|  | ave (-) | std dev | ave (-) | std dev |  |  | ave (-) | std dev | ave (-) | std dev
|  | 0.09997 | 8.8993e-06 | 0.00014787 | 8.4592e-07 | 0.099822 |  | 0.064452 | 8.0314e-06 | 0.00014686 | 1.1057e-06 | 0.064306 | 0.32714
[6 empty] | 0 | [5 empty] | 0 | #DIV/0!
AC |  | ORW-518.RA | [5 empty] | ORW-006.RA
|  | S [A/W] | 0.50782 | [5 empty] | S [A/W] | 0.349 | 0.35 |  | Respons from ORW-518.RA
|  | light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V]
|  | ave | std dev | ave | std dev |  |  | ave | std dev | ave | std dev
Femto on DC setting |  | 4.3633 | 2.3615e-05 | 0.0002603 | 1.4613e-05 | 4.363 | [5 empty] | 0 | 0 | #VALUE!
Femto on AC setting |  | 4.3504 | 1.8543e-05 | 0.00022958 | 1.3643e-05 | 4.3502 | [5 empty] | 0 | 0 | #DIV/0!
DC Femto on DC setting |  | ORW-518.RA | [5 empty] | ORW-522.RA | [7 empty] | ORW-517.RA
|  | S [A/W] | 0.50782 | [9 empty] | Respons from ORW-518.RA [A/W] | Respons from ORW-517.RA [A/W] |  | S [A/W] | 0.50794
|  | light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] |  |  |  | light [V] |  | dark [V] |  | net [V]
|  | ave (-) | std dev | ave (-) | std dev |  |  | ave | std dev | ave | std dev |  |  |  |  | ave (-) | std dev | ave (-) | std dev
| 2 | 0.099819 | 5.1804e-06 | 0.00015288 | 8.6315e-07 | 0.099666 | 1 | 0.068739 | 3.6331e-06 | -0.00014936 | 9.4996e-07 | 0.068889 | 0.351 | 0.35106 | 4 | 0.099822 | 4.3269e-06 | 0.00015029 | 1.3244e-06 | 0.099672
[7 empty] | 3 | 0.068729 | 4.6683e-06 | -0.00015007 | 5.7144e-07 | 0.068879 | 0.35095 | 0.35102
[12 empty] | Average | 0.35098 | 0.35104
[13 empty] | 0.12111 | 0.017671
AC (LIA) |  | ORW-518.RA | [5 empty] | ORW-522.RA | [7 empty] | ORW-517.RA
|  | S [A/W] | 0.50782 | [9 empty] | Respons from ORW-518.RA | Respons from ORW-517.RA |  | S [A/W] | 0.50794
|  | light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] |  |  |  | light [V] |  | dark [V] |  | net [V]
|  | ave | std dev | ave | std dev |  |  | ave | std dev | ave | std dev |  |  |  |  | ave | std dev | ave | std dev
Femto on DC setting | 1 | 4.3522 | 2.064e-05 | 0.00014271 | 1.6315e-05 | 4.3521 | 2 | 3.0076 | 0.00047621 | 0.00019145 | 1.7734e-05 | 3.0074 | 0.35092 | 0.35096 | 3 | 4.3528 | 1.5553e-05 | 0.00018944 | 1.4411e-05 | 4.3526
[6 empty] | 0 | 4 | 3.008 | 1.672e-05 | 0.0001938 | 8.1103e-06 | 3.0078 | 0.35097 | 0.35101
Femto on AC setting | 7 | 4.3403 | 1.7134e-05 | 0.00021157 | 1.4271e-05 | 4.3401 | 5 | 3 | 2.0842e-05 |  |  | 3 | 0.35101
[7 empty] | 6 | 2.9999 | 1.7364e-05 | 0.00021136 | 1.5143e-05 | 2.9997 | 0.35098
Gain correction factor
|  | ORW-518.RA | ORW-522.RA | ORW-517.RA
Ratio |  | 0.43666 | 0.43662 | 0.43669 |  | 2.0145
times 2 for chopper duty cycle |  | 0.87333 | 0.87324 | 0.87338 |  | 2.2901
Duty cycle measurement
ORW-518.RA |  |  |  | Thermopile ORW-506.RA
Chopper freq 31.5 Hz |  | ave | std dev |  | ave | std dev
ORW-518.RA with ORE-111.RA on DC; integration mode NPLC, NPLC: 100; range: auto; input impedance: 10 Mohm; auto zero: on; Null: off; chopper on | AC | -0.049957 | 4.6416e-06 | Chopper on; NPLC: 1 | 8.5795e-05 | 3.8858e-07
ORW-518.RA with ORE-111.RA on DC chopper off | DC | -0.099784 | 1.4461e-06 | Chopper off | 0.00017292 | 3.8838e-07
Ratio |  | 0.50065 |  | Ratio | 0.49615
ORW-518.RA with ORE-111.RA on AC chopper on | AC | 0.049434 | 2.811e-06
ORW-518.RA with ORE-111.RA on DC chopper off | DC | 0.099779 | 7.6253e-06
Ratio |  | 0.49543
𝑠_𝑈𝑈𝑇 (𝜆)=𝑠_𝑆𝑇𝐷 (𝜆) (𝑉_𝑈𝑈𝑇 (𝜆))/(𝑉_𝑆𝑇𝐷 (𝜆))

## 7 Sep
Temperature [deg Celsius] | 23.9
Relative humidity [%] | 52.5
Chopper | 3502 | 2 blade wheel | 37 Hz
Amplifier | ORE-546.RA | 1000 | FBW
Bewteen Femto amp & lock-in, used differential cable
LIA settings | 7214
Input limit
Sensitivity | 100 mV
Time const | 2s
Coupling | AC
Input shell | Float
Input mode | Voltage
Input connector | A-B
Input device | FET
Notch filter | 50 Hz
AC gain | OFF
Ref Source | ExtR RCU
Output slope | 12 dB Oct
Sync time const | OFF
Time const mode | Normal
DVM | ORE-111.RA | ave of 20 readings
laser power controller, HeNe laser | 65 uW
Mathematical model | Spyro = Ssi*Gsi*(Vpyro/Vsi)
[V/W] = A/W*Ω*(V/V)
AC (LIA) @ 37 Hz Chopper Signal Recovery model 197 | ORW-518.RA | [5 empty] | Pyro | [7 empty] | ORW-517.RA
| S [A/W] | 0.50782 | Sensitivity [mV] | 200 |  |  | Sensitivity [mV] | 200 |  |  |  | Respons from ORW-518.RA | Respons from ORW-517.RA [V/W] |  | S [A/W] | 0.50794 | Sensitivity [mV] | 10
| light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] |  |  |  | light [V] |  | dark [V] |  | net [V] | Range to range conversion
| ave | std dev | ave | std dev |  |  | ave | std dev | ave | std dev |  |  |  |  | ave | std dev | ave | std dev
[5 empty] | 0 | 1 | 4.7129 | 0.0013658 |  |  | 4.7129 | #DIV/0! | 445.5116 | 3 | 5.3743 | 0.00064987 | 0.0010097 | 0.00035796 | 5.3733 | 0.26867
[5 empty] | 0 |  | 94.3 mV on LIA | [7 empty] | 5.37 mV on LIA |  |  |  | #VALUE!
[6 empty] | 2 | 4.7255 | 0.0040519 | 0.0063583 | 0.0011761 | 4.7192 | #DIV/0! | 8922.0789
94.6 mV on LIA
AC (LIA) @ 22 Hz Chopper Signal Recovery model 197 | ORW-518.RA | [5 empty] | Pyro | [7 empty] | ORW-517.RA
| S [A/W] | 0.50782 | Sensitivity [mV] | 200 |  |  | Sensitivity [mV] | 500 |  |  |  | Respons from ORW-518.RA | Respons from ORW-517.RA [V/W] |  | S [A/W] | 0.50794 | Sensitivity [mV] | 500
| light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] |  |  |  | light [V] |  | dark [V] |  | net [V] | Range to range conversion
| ave | std dev | ave | std dev |  |  | ave [V] on LIA | std dev | ave | std dev |  |  |  |  | ave [V] on LIA | std dev | ave | std dev
[5 empty] | 0 | 1 | 0.1972 |  |  |  | 0.1972 | #DIV/0! | 14107.8546 | 2 | 0.0071 |  |  |  | 0.0071 | 0.0071
[5 empty] | 0 | [13 empty] | 0
AC (LIA) @ 17 Hz Chopper Signal Recovery model 197 | ORW-518.RA | [5 empty] | Pyro | [7 empty] | ORW-517.RA
| S [A/W] | 0.50782 | Sensitivity [mV] | 200 |  |  | Sensitivity [mV] | 500 |  |  |  | Respons from ORW-518.RA | Respons from ORW-517.RA [V/W] |  | S [A/W] | 0.50794 | Sensitivity [mV] | 500
| light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] |  |  |  | light [V] |  | dark [V] |  | net [V] | Range to range conversion
| ave | std dev | ave | std dev |  |  | ave [V] on LIA | std dev | ave | std dev |  |  |  |  | ave [V] on LIA | std dev | ave | std dev
[5 empty] | 0 | 1 | 0.246 |  |  |  | 0.246 | #DIV/0! | 17354.6167 | 2 | 0.0072 |  |  |  | 0.0072 | 0.0072
[5 empty] | 0 | [13 empty] | 0
AC (LIA) @ 13 Hz Chopper Newport model 3502 | ORW-518.RA | [5 empty] | Pyro | [7 empty] | ORW-517.RA
| S [A/W] | 0.50782 | Sensitivity [mV] | 200 |  |  | Sensitivity [mV] | 500 |  |  |  | Respons from ORW-518.RA | Respons from ORW-517.RA [V/W] |  | S [A/W] | 0.50794 | Sensitivity [mV] | 500
| light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] |  |  |  | light [V] |  | dark [V] |  | net [V] | Range to range conversion
| ave | std dev | ave | std dev |  |  | ave [V] on LIA | std dev | ave | std dev |  |  |  |  | ave [V] on LIA | std dev | ave | std dev
[5 empty] | 0 | 1 | 0.2993 |  |  |  | 0.2993 | #DIV/0! | 20544.1138 | 2 | 0.0074 |  |  |  | 0.0074 | 0.0074
[5 empty] | 0 | [13 empty] | 0
AC (LIA) @ 5 Hz Chopper Newport model 3502 | ORW-518.RA | [5 empty] | Pyro | [7 empty] | ORW-517.RA
| S [A/W] | 0.50782 | Sensitivity [mV] | 200 |  |  | Sensitivity [mV] | 1000 |  |  |  | Respons from ORW-518.RA | Respons from ORW-517.RA [V/W] |  | S [A/W] | 0.50794 | Sensitivity [mV] | 1000
| light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] |  |  |  | light [V] |  | dark [V] |  | net [V] | Duty cycle
| ave | std dev | ave | std dev |  |  | ave [V] on LIA | std dev | ave | std dev |  |  |  |  | ave [V] on LIA | std dev | ave | std dev |  | 0.49543
[5 empty] | 0 | 1 | 0.486 |  |  |  | 0.486 | #DIV/0! | 33816.2795 | 2 | 0.0073 |  |  |  | 0.0073
[5 empty] | 0 | [7 empty] | 68255.8619 | [5 empty] | 0

## 10 Sep
Temperature [deg Celsius] |  | 22.3 | 22.3
Relative humidity [%] |  | 54.4 | 57.4
Chopper
Amplifier |  | ORE-546.RA | 952.65 | FBW
Bewteen Femto amp & lock-in, used differential cable
LIA settings |  | 7214
Input limit
Sensitivity |  | 100 mV
Time const |  | 5s
Coupling |  | AC
Input shell |  | Float
Input mode |  | Voltage
Input connector |  | A-B
Input device |  | FET
Notch filter |  | 50 Hz
AC gain |  | OFF
Ref Source |  | ExtR RCU
Output slope |  | 12 dB Oct
Sync time const |  | OFF
Time const mode |  | Normal
DVM |  | ORE-111.RA | ave of 20 readings
laser power controller, HeNe laser |  | 208 uW
Mathematical model |  | Spyro = Ssi*Gsi*(Vpyro/Vsi)
[V/W] = A/W*Ω*(V/V)
AC (LIA) @ 135 Hz Chopper model 197 (ORE-162.RA) 208 uW |  | ORW-518.RA | [5 empty] | Pyro | [7 empty] | ORW-517.RA | [6 empty] | ORW-522.RA
|  | S [A/W] | 0.50782 | Sensitivity [mV] | 200 |  |  | Sensitivity [mV] | 200 |  |  |  | Respons from ORW-518.RA | Respons from ORW-517.RA [V/W] |  | S [A/W] | 0.50794 | Sensitivity [mV] | 200 | [8 empty] | Respons from ORW-518.RA | Respons from ORW-517.RA
|  | light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] |  |  |  | light [V] |  | dark [V] |  | net [V] | Duty cycle |  | light [V] |  | dark [V] |  | net [V]
|  | ave | std dev | ave | std dev |  |  | ave [V] on LIA | std dev | ave | std dev |  |  |  |  | ave [V] on LIA | std dev | ave | std dev |  | 0.49543 |  | ave | std dev | ave | std dev
| 3 | 0.0301 |  |  |  | 0.0301 | 1 | 0.136 |  |  |  | 0.136 | 2185.826 | 2186.3425 | 2 | 0.0301 |  |  |  | 0.0301 |  | 4 | 0.0208 |  |  |  | 0.0208 | 0.35092 | 0.351
[6 empty] | 0 | [7 empty] | 1083.1886 | [5 empty] | 0
[8 empty] | Pyro | [7 empty] | ORW-517.RA
[8 empty] | Sensitivity [mV] | 200 |  |  |  | Respons from ORW-518.RA | Respons from ORW-517.RA [V/W] |  | S [A/W] | 0.50794 | Sensitivity [mV] | 100
[8 empty] | light [V] |  | dark [V] |  | net [V] |  |  |  | light [V] |  | dark [V] |  | net [V] | Duty cycle
[8 empty] | ave [V] on LIA | std dev | ave | std dev |  |  |  |  | ave [V] on LIA | std dev | ave | std dev |  | 0.49543
[7 empty] | 5 | 0.136 |  |  |  | 0.136 | #DIV/0! | 2186.3425 | 6 | 0.0301 |  |  |  | 0.0301
[8 empty] | 0.00068 | [5 empty] | 1083.1886 |  | 0.000301 |  |  |  | 0.000301 | 0.0001505
AC (LIA) @ 135 Hz Chopper model 197 (ORE-162.RA) 208 uW femto in AC mode |  | ORW-518.RA | [5 empty] | Pyro | [7 empty] | ORW-517.RA | [6 empty] | ORW-522.RA
|  | S [A/W] | 0.50782 | Sensitivity [mV] | 100 |  |  | Sensitivity [mV] | 200 |  |  |  | Respons from ORW-518.RA | Respons from ORW-517.RA [V/W] |  | S [A/W] | 0.50794 | Sensitivity [mV] | 100 |  |  |  | Sensitivity [mV] | 100 |  |  |  | Respons from ORW-518.RA | Respons from ORW-517.RA
|  | light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] |  |  |  | light [V] |  | dark [V] |  | net [V] | Duty cycle |  | light [V] |  | dark [V] |  | net [V]
|  | ave | std dev | ave | std dev |  |  | ave [V] on LIA | std dev | ave | std dev |  |  |  |  | ave [V] on LIA | std dev | ave | std dev |  | 0.49862 |  | ave | std dev | ave | std dev
| 4 | 3.0088 | 1.9637e-05 | 7.1255e-05 | 1.3058e-05 | 1.5044 | 1 | 6.785 | 0.0045004 | 0.002577 | 0.00043251 | 6.7824 | 1087.5185 | 1087.7655 | 2 | 3.0089 | 1.986e-05 | 7.2603e-05 | 1.8163e-05 | 1.5044 |  | 8 | 2.0802 | 1.8327e-05 | 0.00010261 | 1.3447e-05 | 2.0801 | 0.3501 | 0.35019
| 9 | 3.0102 | 2.9849e-05 | 5.3441e-05 | 1.8128e-05 | 1.5051 | 3 | 6.784 | 0.0033621 | 0.0021602 | 0.00036785 | 6.7819 | 1087.4337 | 1087.6807 | 6 | 3.0102 | 2.7254e-05 | 0.00011698 | 1.4068e-05 | 1.505
[7 empty] | 5 | 6.7791 | 0.0044504 | 0.0013228 | 0.00048242 | 6.7778 | 1086.7791 | 1086.5634
[7 empty] | 7 | 6.7779 | 5.283775-3 | 0.001441 | 0.00049323 | 6.7765 | 1086.0867 | 1086.3543
1087.0227
0.65385
0.06015
Duty cycle measurement
Thermopile ORW-506.RA |  |  | light |  | dark |  |  | net [V]
Chopper freq 135 Hz; power 298 uW |  |  | ave [V] | std dev [V] | ave [V] | std dev [V]
ORE-111.RA on AC chopper on | AC | Chopper on | 8.2422e-05 | 8.0307e-07 | -1.1266e-06 | 7.3228e-07 |  | 8.3548e-05
ORE-111.RA on DC chopper off | DC | Chopper off | 0.00016548 | 7.5716e-07 | -2.0768e-06 | 6.5871e-07 |  | 0.00016756
[6 empty] | Ratio |  | 0.49862

## 13 Sep
Temperature [deg Celsius] |  | 22.9
Relative humidity [%] |  | 54.4
Chopper |  | Newport 3502 | ORE-070.RA
FemtoAmplifier, used only for photodiode |  | ORE-546.RA | 1000 | FBW
Bewteen Femto amp & lock-in, used differential cable
LIA settings Model 7214
Input limit
Sensitivity |  | 100 mV
Time const |  | 5s
Coupling |  | AC
Input shell |  | Float | [13 empty] | ORW-522.RA | [6 empty] | ORW-522.RA
Input mode |  | Voltage | [10 empty] | Respons from ORW-518.RA | Respons from ORW-517.RA [V/W] |  | S [A/W] | 0.50794 | Sensitivity [mV] | 200
Input connector |  | A-B | [9 empty] | net [V] |  |  |  | light [V] |  | dark [V] |  | net [V] | Duty cycle |  | light [V] |  | dark [V]
Input device |  | FET
Notch filter |  | 50 Hz | [13 empty] | ave [V] on LIA | std dev | ave | std dev |  | 0.49543 |  | ave | std dev | ave
AC gain |  | OFF | [9 empty] | 0.136 | 2294.4691 | 2295.0113 | 2 | 0.0301 |  |  |  | 0.0301 |  | 4 | 0.0208
Ref Source |  | ExtR RCU | [11 empty] | 1137.0268 | [5 empty] | 0
Output slope |  | 12 dB Oct | [13 empty] | ORW-517.RA
Sync time const |  | OFF | [10 empty] | Respons from ORW-518.RA | Respons from ORW-517.RA [V/W] |  | S [A/W] | 0.50794 | Sensitivity [mV] | 100
Time const mode |  | Normal | [9 empty] | net [V] |  |  |  | light [V] |  | dark [V] |  | net [V] | Duty cycle
DVM |  | ORE-111.RA | ave of 20 readings
laser power controller, HeNe laser |  | 208 uW | [8 empty] | ORW-522.RA |  |  |  |  | ave [V] on LIA | std dev | ave | std dev |  | 0.49543
Mathematical model |  | Spyro = Ssi*Gsi*(Vpyro/Vsi) | [9 empty] | 0.136 | #DIV/0! | 2295.0113 | 6 | 0.0301 |  |  |  | 0.0301
|  | [V/W] = A/W*Ω*(V/V) | [11 empty] | 1137.0268 |  | 0.000301 |  |  |  | 0.000301 | 0.0001505
AC (LIA) @ 140 Hz Chopper model 197 (ORE-162.RA) 208 uW |  | ORW-522.RA ORW-518.RA | [5 empty] | Pyro | [7 empty] | ORW-517.RA | [6 empty] | ORW-522.RA
|  | S [A/W] | 0.50782 | Sensitivity [mV] | 200 |  |  | Sensitivity [mV] | 200 |  |  |  | Respons from ORW-518.RA | Respons from ORW-517.RA [V/W] |  | S [A/W] | 0.50794 | Sensitivity [mV] | 100 |  |  |  | Sensitivity [mV] | 100 |  |  |  | Respons from ORW-518.RA | Respons from ORW-517.RA
|  | light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] |  | net [V] |  |  |  | light [V] |  | dark [V] |  | net [V] | Duty cycle |  | light [V] |  | dark [V] |  | net [V]
|  | ave | std dev | ave | std dev |  |  | ave [V] on LIA | std dev | ave | std dev |  |  |  |  | ave [V] on LIA | std dev | ave | std dev |  | 0.49862 |  | ave | std dev | ave | std dev
| 3 | 0.0301 |  |  |  | 0.0301 | 1 | 0.136 |  |  |  | 6.7824 | 1141.572 | 1141.8312 | 2 | 3.0089 | 1.986e-05 | 7.2603e-05 | 1.8163e-05 | 1.5044 |  | 8 | 2.0802 | 1.8327e-05 | 0.00010261 |  | 0.0208 | 0.35092 | 0.351
[12 empty] | 6.7819 | 1141.4829 | 1141.7422 | 6 | 3.0102 | 2.7254e-05 | 0.00011698 | 1.4068e-05 | 1.505
[8 empty] | Pyro |  |  |  | 6.7778 | 1140.7958 | 1140.5694
[8 empty] | Sensitivity [mV] | 200 |  |  | 6.7765 | 1140.0689 | 1140.3499
[8 empty] | light [V] |  | dark [V] |  |  |  | 1141.0515
0.68635
[8 empty] | ave [V] on LIA | std dev | ave | std dev |  |  | 0.06015
[7 empty] | 5 | 0.136
0.00068
AC (LIA) @ 135 Hz Chopper model 197 (ORE-162.RA) 208 uW femto in AC mode |  | ORW-518.RA | [5 empty] | Pyro
|  | S [A/W] | 0.50782 | Sensitivity [mV] | 100 |  |  | Sensitivity [mV] | 200 | [18 empty] | Respons from ORW-518.RA | Respons from ORW-517.RA
|  | light [V] |  | dark [V] |  | net [V] |  | light [V] |  | dark [V] | [16 empty] | net [V]
|  | ave | std dev | ave | std dev |  |  | ave [V] on LIA | std dev | ave | std dev | [14 empty] | std dev
| 4 | 3.0088 | 1.9637e-05 | 7.1255e-05 | 1.3058e-05 | 1.5044 | 1 | 6.785 | 0.0045004 | 0.002577 | 0.00043251 | [14 empty] | 1.3447e-05 | 2.0801 | 0.3501 | 0.35019
| 9 | 3.0102 | 2.9849e-05 | 5.3441e-05 | 1.8128e-05 | 1.5051 | 3 | 6.784 | 0.0033621 | 0.0021602 | 0.00036785
[7 empty] | 5 | 6.7791 | 0.0044504 | 0.0013228 | 0.00048242
[7 empty] | 7 | 6.7779 | 5.283775-3 | 0.001441 | 0.00049323
0.35055
Duty cycle measurement
Thermopile ORW-506.RA |  |  | light |  | dark |  |  | net [V]
Chopper freq 135 Hz; power 298 uW |  |  | ave [V] | std dev [V] | ave [V] | std dev [V]
ORE-111.RA on AC chopper on | AC | Chopper on | 8.2422e-05 | 8.0307e-07 | -1.1266e-06 | 7.3228e-07 |  | 8.3548e-05
ORE-111.RA on DC chopper off | DC | Chopper off | 0.00016548 | 7.5716e-07 | -2.0768e-06 | 6.5871e-07 |  | 0.00016756
[6 empty] | Ratio |  | 0.49862

## 14 Sep
Source: FLEXPOINT FP-D-905-5P-C-C Dot laser Applied voltage: 5,5 V |  | Experimental description: |  |  |  | Monitor photodiode (DC signal) Amplify with FEMTO (NMI22217); DC; FBW; Gain = 10^6 |  |  |  | Pyro absorption at 633nm: 0,964 at 905 nm: 0,954 |  | Temp: 22,7 degC |  | Hum: 22,2
|  | ORW-522.RA | [6 empty] | Pyro (S/N 505439) | [6 empty] | ORW-522.RA |  |  |  |  | ORW-522.RA
|  | Femto amp (ORE-546.RA); FBW; AC | Signal Recovery LIA 7124 (ORE-164.RA) | Coordinates (357,5;52) | [5 empty] | Signal Recovery LIA 7124 (ORE-164.RA) | Coordinates (357,5;52) |  |  |  |  | Femto amp (ORE-546.RA); FBW; AC | Signal Recovery LIA 7124 (ORE-164.RA) | Coordinates (357,5;52)
|  | Light |  | Dark |  |  |  |  | Light |  | Dark |  |  |  |  | Light |  | Dark
LIA SENS | FEMTO Gain | Ave | sdev | Ave | sdev | Nett | LIA SENS[mV] | FEMTO Gain | Ave[mV] | sdev | Ave[mV] | sdev | Nett[V] | LIA SENS | FEMTO Gain | Ave | sdev | Ave | sdev | Nett | ORW-522.RA Ave[V] | SENS Corr | Pyro abs corr | Chopper Duty Cycle | Pyro resp[V/W] | ORW-522.RA resp[A/W]
500 | 952.65 | 478.1609 | 0.26907 | -0.084463 | 0.016106 | 478.2454 | 500 | 0 | 2016.66 | 2.178 | -0.082032 | 0.015 | 2.0167 | 500 | 952.65 | 478.0126 | 0.14101 | -0.097934 | 0.012782 | 478.1105 | 0.47818 | 1 | 0.98963 | 0.49543 | 1087 | 0.54041 | 0.48654
500 | 952.65 | 536.7624 | 0.23951 | -0.23475 | 0.014503 | 536.9972 | 500 | 0 | 2247.792 | 2.9891 | -0.23712 | 0.017094 | 2.248 | 500 | 952.65 | 537.7536 | 0.33675 | -0.24092 | 0.012487 | 537.9945 | 0.5375 | 1 | 0.98963 | 0.49543 | 1087 | 0.54495 | 0.49063
500 | 952.65 | 533.742 | 0.33813 | -0.26734 | 0.013093 | 534.0093 | 500 | 0 | 2244.336 | 2.1923 | -0.28545 | 0.015408 | 2.2446 | 500 | 952.65 | 537.3128 | 0.2381 | -0.28408 | 0.015079 | 537.5969 | 0.5358 | 1 | 0.98963 | 0.49543 | 1087 | 0.54406 | 0.48983
500 | 952.65 | 538.9825 | 0.26349 | -0.30701 | 0.011351 | 539.2895 | 500 | 0 | 2263.126 | 1.2022 | -0.29384 | 0.013507 | 2.2634 | 500 | 952.65 | 540.4352 | 0.37493 | -0.26657 | 0.015142 | 540.7018 | 0.54 | 1 | 0.98963 | 0.49543 | 1087 | 0.54376 | 0.48956 | polarization changed by 90 deg
2018-09-25 | [8 empty] | 633 nm
|  | Pyro | [6 empty] | Si-trap | [6 empty] | Pyro
|  | ORS-003.RA | [6 empty] | ORW-517.RA | [6 empty] | ORS-003.RA |  |  |  |  | ORS-003.RA
|  | Femto amp (ORE-546.RA); FBW; AC | Signal Recovery LIA 7124 (ORE-164.RA) | Coordinates (194;16,75) | [5 empty] | Signal Recovery LIA 7124 (ORE-164.RA) | Coordinates (306;47) |  |  |  |  | Femto amp (ORE-546.RA); FBW; AC | Signal Recovery LIA 7124 (ORE-164.RA) | Coordinates (194;16,75)
|  | Light |  | Dark |  |  |  |  | Light |  | Dark |  |  |  |  | Light |  | Dark
LIA SENS | FEMTO Gain | Ave | sdev | Ave | sdev | Nett | LIA SENS[mV] | FEMTO Gain | Ave[mV] | sdev | Ave[mV] | sdev | Nett[V] | LIA SENS | FEMTO Gain | Ave | sdev | Ave | sdev | Nett | ORS-003.RA Ave[V] | SENS Corr | Pyro abs corr | Chopper Duty Cycle | ORW-517 resp[A/W] | ORS-003.RA resp[A/W]
100 | 1171341666.6667 | 7.1243 | 0.049463 | 0.00015724 | 1.4526e-05 | 7.1242 | 100 | 949.7387 | 3.5024 | 0.018295 | 0.000117 | 1.6009e-05 | 3.5023 | 100 | 1171341666.6667 | 7.1846 | 0.011501 | 0.00012598 | 1.5698e-05 | 7.1845 | 7.1543 | 1 | 1 | 0.49543 | 0.50794 | 4.1681e-07 | 35 Hz
100 | 1171341666.6667 | 7.2302 | 3.4062e-05 | 0.00011955 | 1.6666e-05 | 7.2301 | 100 | 949.7387 | 3.5488 | 0.00046786 | 0.00011203 | 1.6953e-05 | 3.5487 | 100 | 1171341666.6667 | 7.2298 | 0.00054276 | 9.7366e-05 | 1.2523e-05 | 7.2297 | 7.2299 | 1 | 1 | 0.49543 | 0.50794 | 4.157e-07 | 75 Hz
100 | 1171341666.6667 | [5 empty] | 100 | 949.7387 | [5 empty] | 100 | 1171341666.6667 | [5 empty] | #DIV/0! | 1 | 1 | 0.49543 | 0.50794 | #DIV/0!
100 | 1171341666.6667 | [5 empty] | 100 | 949.7387 | [5 empty] | 100 | 1171341666.6667 | [5 empty] | #DIV/0! | 1 | 1 | 0.49543 | 0.50794 | #DIV/0!
|  | Si-photodiode | [6 empty] | Pyro | [6 empty] | Si-photodiode
|  | ORW-522.RA | [6 empty] | ORS-003.RA | [6 empty] | ORW-522.RA |  |  |  |  | ORW-522.RA
|  | Femto amp (ORE-546.RA); FBW; AC | Signal Recovery LIA 7124 (ORE-164.RA) | Coordinates (495;50) | [5 empty] | Signal Recovery LIA 7124 (ORE-164.RA) | Coordinates (194;16,75) |  |  |  |  | Femto amp (ORE-546.RA); FBW; AC | Signal Recovery LIA 7124 (ORE-164.RA) | Coordinates (495;50)
|  | Light |  | Dark |  |  |  |  | Light |  | Dark |  |  |  |  | Light |  | Dark
LIA SENS | FEMTO Gain | Ave | sdev | Ave | sdev | Nett | LIA SENS[mV] | FEMTO Gain | Ave[mV] | sdev | Ave[mV] | sdev | Nett[V] | LIA SENS | FEMTO Gain | Ave | sdev | Ave | sdev | Nett | ORW-522.RA Ave[V] | SENS Corr | Pyro abs corr | Chopper Duty Cycle | ORS-003 resp[A/W] | ORW-522.RA resp[A/W]
100 | 949.7387 | 2.4526 | 0.00031686 | 8.8894e-05 | 1.6201e-05 | 2.4525 | 100 | 1171341666.6667 | 7.2341 | 2.0614e-05 | 0.00010204 | 1.4153e-05 | 7.234 | 100 | 949.7387 | 2.4525 | 1.1433e-05 | 8.8208e-05 | 1.1182e-05 | 2.4524 | 2.4525 | 1 | 1 | 0.49543 | 4.157e-07 | 0.35084 | 75 Hz
100 | 949.7387 |  |  |  |  | 0 | 100 | 1171341666.6667 |  |  |  |  | 0 | 100 | 949.7387 |  |  |  |  | 0 | 0 | 1 | 1 | 0.49543 | 4.157e-07 | #DIV/0! | 75 Hz
100 | 949.7387 |  |  |  |  | 0 | 100 | 1171341666.6667 |  |  |  |  | 0 | 100 | 949.7387 |  |  |  |  | 0 | 0 | 1 | 1 | 0.49543 | 4.157e-07 | #DIV/0!
100 | 949.7387 |  |  |  |  | 0 | 100 | 1171341666.6667 |  |  |  |  | 0 | 100 | 949.7387 |  |  |  |  | 0 | 0 | 1 | 1 | 0.49543 | 4.157e-07 | #DIV/0!
Chopper duty cycle
[24 empty] | Av | Sdev
| 2018-09-25 | [8 empty] | 905 nm |  | n = 75 HZ | [10 empty] | chopper | 0.49298 | 0.00035148
[23 empty] | no chopper | 0.93156 | 0.00037107
|  | ORW-522.RA | [6 empty] | Pyro (S/N 505439) |  |  | ORS-003 |  |  |  | ORW-522.RA |  |  |  |  | ORW-522.RA |  | Duty cycle: | 0.5292
|  | Femto amp (ORE-546.RA); FBW; AC | Signal Recovery LIA 7124 (ORE-164.RA) | Coordinates (357,5;52) | [5 empty] | Signal Recovery LIA 7124 (ORE-164.RA) | Coordinates (108;24,5) | Coordinates (58;19,5) |  |  |  | Femto amp (ORE-546.RA); FBW; AC | Signal Recovery LIA 7124 (ORE-164.RA) | Coordinates (357,5;52)
|  | Light |  | Dark |  |  |  |  | Light |  | Dark |  |  |  |  | Light |  | Dark
LIA SENS[V] | FEMTO Gain | Ave | sdev | Ave | sdev | Nett | LIA SENS[mV] | FEMTO Gain | Ave[mV] | sdev | Ave[mV] | sdev | Nett[V] | LIA SENS | FEMTO Gain | Ave | sdev | Ave | sdev | Nett | ORW-522.RA Ave[V] | SENS Corr | Pyro abs corr | Chopper Duty Cycle | Pyro resp[V/W] | ORW-522.RA resp[A/W]
0.2 | 949.7387 | 3.7647 | 1.6981e-05 | 0.00013447 | 1.6371e-05 | 3.7646 | 1 | 0 | 6.766 | 0.00087241 | 0.00013844 | 2.036e-05 | 6.7659 | 0.2 | 949.7387 | 3.7657 | 1.3372e-05 | 8.6729e-05 | 1.417e-05 | 3.7656 | 3.7651 | 0.2 | 0.98963 | 0.52977 | 1087 | 0.23795 | The pyro resp was determined at 135 Hz
0.2 | 949.7387 | 3.773 | 1.7133e-05 | 9.0566e-05 | 1.3101e-05 | 3.7729 | 1 | 0 | 3.5224 | 0.0022019 | 0.00028677 | 0.00035769 | 3.5222 | 0.2 | 949.7387 | 3.771 | 1.6529e-05 | 0.00010786 | 1.6287e-05 | 3.7709 | 3.7719 | 0.2 | 0.98963 | 0.49897 | 1087 | 0.48619 | n = 135 HZ
0.2 | 949.7387 | 3.7742 | 1.7814e-05 | 0.00012028 | 1.4589e-05 | 3.774 | 0.2 | 1171341666.6667 | 7.2137 | 0.0004379 | 0.00013145 | 1.4215e-05 | 7.2136 | 0.2 | 949.7387 | 3.7732 | 0.00054898 | 0.00011861 | 1.4189e-05 | 3.7731 | 3.7735 | 1 | 1 | 0.49817 | 4.157e-07 | 0.53837 | ORS-003.RA 75 Hz
0.2 | 949.7387 |  |  |  |  | 0 | 1 | 0 |  |  |  |  | 0 | 0.2 | 949.7387 |  |  |  |  | 0 | 0 | 0.2 | 0.98963 | 0.49817 | 1087 | #DIV/0! | n = 135 HZ
0.52914
[7 empty] | Chopper duty cycle for 7/5 blade Newport 3502 chopper | [6 empty] | Chopper duty cycle for 7/5 blade Newport 3502 chopper - use 5 slot
Measured with ORW-506.RA
[7 empty] | Chopper |  | No chopper |  | Duty cycle |  |  | Chopper |  |  |  | No Chopper |  |  |  | Duty cycle
[7 empty] | Ave | sdev | Ave | sdev |  |  |  | Ave | dark | sdev | Nett | Ave | dark | sdev | Nett
[7 empty] | 0.49298 | 0.00035148 | 0.93156 | 0.00037107 | 0.5292 |  |  | 3.3909 | 0.0015176 | 0.0015488 | 3.3894 | 6.7954 | 0.0019034 | 0.0032715 | 6.7935 | 0.49891 | Measured with Gentec pyro and 2nd Newport 3502 chopper with 42/30 blade - use 30 slot - 592,866 Hz
[7 empty] | 0.49467 | 0.00022337 | 0.93306 | 0.00075688 | 0.53016 |  |  | 3.3939 | 0.0020378 | 0.0021248 | 3.3919 | 6.7987 | 0.0017652 | 0.0047539 | 6.797 | 0.49903 | Measured with Gentec pyro and 2nd Newport 3502 chopper with 42/30 blade - use 30 slot - 592,866 Hz
[7 empty] | 0.49388 | 0.00033873 | 0.93346 | 0.00045105 | 0.52909 |  |  | 3.5016 | 0.00028764 | 0.00049316 | 3.5013 | 7.0414 | 0.00028977 | 0.0012345 | 7.0411 | 0.49727 | Measured with ORS-003.RA pyro and 2nd Newport 3502 chopper with 42/30 blade - use 30 slot - 592,866 Hz
[7 empty] | 0.49358 | 0.0003804 | 0.9302 | 0.00061946 | 0.53062 |  |  | 3.503 | 0.00030923 | 0.0002382 | 3.5027 | 7.0411 | 0.00030361 | 0.001197 | 7.0408 | 0.49748 | Measured with ORS-003.RA pyro and 2nd Newport 3502 chopper with 42/30 blade - use 30 slot - 592,866 Hz
[11 empty] | 0.52977 | [10 empty] | 0.49817 | Average of all measurements
[11 empty] | do not use this value | [10 empty] | 0.49897 | Average of measurements done with Gentec pyro

## 22 Nov
Aperture diameter [mm]
|  |  | 7.96 |  | UPD-50-UD
|  |  | 8.02 |  | Femto gain: 10^6; DC; L; FBW [Ohm] | Light [mV] | SDEV [uV] | Dark [mV] | SDEV [uV] | Nett [mV]
|  |  | 8.02 |  | 996687.3 | 44.5389 | 64.3108 | 34.7721 | 54.631 | 9.7668
|  |  | 8 |  |  | 45.3872 | 41.0886 | 35.6295 | 49.1096 | 9.7577
|  |  | 8 | [5 empty] | Average [mV] | 9.7623
|  |  | 7.98 |  | ORW-522.RA
|  |  | 8.02 |  | Femto gain: 10^3; DC; L; FBW [Ohm] | Light [mV] | SDEV [uV] | Dark [mV] | SDEV [uV] | Nett [mV]
|  |  | 8 |  | 952.632 | 113.5248 | 19.5891 | -0.00016751 | 0.6872 | 113.525
|  |  | 7.98 | [5 empty] | Current [A] | 0.00011917
8
|  | Average | 7.998 |  | ORW-522.RA Power responsivity at 905 nm [A/W] | 0.507 |  |  | Power measured with ORW-522.RA [W] | 0.00023505
|  | SDEV | 0.019889 |  | Irradiance measured with ORW-522.RA at 905 nm [W/m2] | 4.6785
Aperture area
|  |  | 50.2404 | [mm^2]
|  |  | 5.024e-05 | [m^2] | Irradiance responsivity of UPD-50-UD [mV/W.m^2] | 2.0866
Nuutste resultate
| Aperture diam on ORW522 [mm] | 2.7 |  |  | LIA: SENS = 50mV
UPD-50-UD
[5 empty] | Femto gain: 10^8; DC; L; FBW [Ohm] | Light [V] | SDEV [uV] | Dark [V] | SDEV [uV] | Nett [V]
[5 empty] | 100000000 | 3.9128 |  | 0.0047039 |  | 3.9081
[6 empty] | 3.9888 |  | 0.0031018 |  | 3.9857
[9 empty] | Average [V] | 3.9469
ORW-522.RA
[5 empty] | Femto gain: 10^5; DC; L; FBW [Ohm] | Light [V] | SDEV [uV] | Dark [V] | SDEV [uV] | Nett [V]
[5 empty] | 100000 | 4.0841 | 19.5891 | 0.00054759 | 0.6872 | 4.0835
[9 empty] | Current [A] | 4.0835e-05
[5 empty] | ORW-522.RA Power responsivity at 905 nm [A/W] | 0.507 |  |  | Power measured with ORW-522.RA [W] | 8.0543e-05
[5 empty] | Irradiance measured with ORW-522.RA at 905 nm [W/m2] | 14.0673 |  |  |  |  | ORW-522.RA Irradiance responsivity at 905 nm [A/W.m-2] | 2.9029e-06
| Aperture area [m^2] |  | 5.7256e-06 | [7 empty] | Irradiance responsivity of UPD-50-UD [A/W.m^2] | 2.8058e-09
[11 empty] | Irradiance responsivity of UPD-50-UD [V/W.m^2] | 0.28058
[11 empty] | Irradiance responsivity of UPD-50-UD [mV/W.m^2] | 280.5758
Calibration of Saab UUT against UPD-50-UD
[5 empty] | Transmittance of ND filter | 0.0015845
UPD-50-UD
[6 empty] | Light [mV] | SDEV [uV] | Dark [mV] | SDEV [uV] | Nett [mV]
[6 empty] | 3.3 |  | 0 |  | 3.3
[9 empty] | Average [mV] | 3.3
SAAB UUT
[6 empty] | Light [mV] | SDEV [uV] | Dark [mV] | SDEV [uV] | Nett [mV]
[6 empty] | 3800 |  | 0 |  | 3800
0
[9 empty] | Average [mV] | 3800
[12 empty] | 323087.2427 | 323.0872

## 7 Jan 2019
Nuutste resultate
| Aperture diam on ORW522 [mm] | 2.7 |  | LIA: SENS = 50mV
| Aperture radius on ORW-522 [m] | 0.00135 |  | IG (Labsphere InGaAs from FO)
|  |  |  | Femto gain: 10^5; DC; L; FBW [Ohm] | Light [V] | SDEV [uV] | Dark [V] | SDEV [uV] | Nett [V]
|  |  |  | 100000 | 4.4553 |  | 0.00055392 |  | 4.4547
[5 empty] | 4.4371 |  | 0.00053368 |  | 4.4366
[8 empty] | Average [V] | 4.4457 | Current [A] | 4.4457e-05
ORW-522.RA
|  |  |  | Femto gain: 10^5; DC; L; FBW [Ohm] | Light [V] | SDEV [uV] | Dark [V] | SDEV [uV] | Nett [V]
|  |  |  | 100000 | 3.8532 | 19.5891 | 0.00054473 | 0.6872 | 3.8527
[8 empty] | Current [A] | 3.8527e-05
|  |  |  | ORW-522.RA Power responsivity at 905 nm [A/W] | 0.507 |  |  | Power | 7.5989e-05 | Irradiance | 13.272
[10 empty] | ORW-522.RA Irradiance responsivity at 905 nm [A/W.m-2] | 2.9029e-06
| Aperture area [m^2] |  | 5.7256e-06 | [6 empty] | Irradiance responsivity of IG [A/W.m^2] | 3.3497e-06
[10 empty] | Irradiance responsivity of IG [V/W.m^2] | 0.33497
[10 empty] | Irradiance responsivity of IG [mV/W.m^2] | 334.9651
Calibration of Saab UUT against IG
|  |  |  | Transmittance of ND filter | 14.66 | 0.1466
IG
[5 empty] | Light [mV] | SDEV [uV] | Dark [mV] | SDEV [uV] | Nett [mV]
[5 empty] | 30 |  | 0 |  | 30
[8 empty] | Average [mV] | 30
SAAB UUT
[5 empty] | Light [mV] | SDEV [uV] | Dark [mV] | SDEV [uV] | Nett [mV]
[5 empty] | 1500 |  | 0 |  | 1500
0
[8 empty] | Average [mV] | 1500
114244.589

## 16 - 21 Jan 2019
Irradiance calibration of SAAB's FEMTO unit (HCA-S-200M-SI) against ORW-522.RA
Date of calibration | 2019-01-16
Wavelength of calibration | 905 nm
Source | Monochromator (600 g/mm --dispersion = 4nm / mm) (ORE-046.RA) + QTH (ORE-080.RA); bandwidth = 10 nm (2,5 mm slitwidth)
HCA-S-200M-SI Serial number | ??????
DVM (ORE-555.RA)
X-Y stage to position detector units (ORE-143.RA) | To ensure exposure to same irradiance field
Nuutste resultate
Aperture diameter on ORW-522.RA
| 2.68 |  | 5.641e-06
| 2.68 |  | 5.641e-06
| 2.66 |  | 5.5572e-06
| 2.7 |  | 5.7256e-06
| 2.7 |  | 5.7256e-06
| 2.6 |  | 5.3093e-06
| 2.72 |  | 5.8107e-06
| 2.7 |  | 5.7256e-06
| 2.7 |  | 5.7256e-06
| 2.7 |  | 5.7256e-06
Average diameter [mm] | 2.684 |  | 5.6587e-06
|  | Exp std dev | 1.4082e-07
|  | ESDM | 4.453e-08
|  |  |  | Chopper frequency (ORE-071.RA) | 75 Hz | (5/7 chopper blade - inner ring)
Aperture diam on ORW522 [mm] | 2.7 |  |  | LIA (ORE-545.RA): SENS = 10mV
Aperture radius on ORW-522 [m] | 0.00135 |  |  | Femto optical receiver (HCA-S-200M-SI (HCA))
[5 empty] | Light [V] | SDEV [V] | Dark [V] | SDEV [V] | Nett [V]
|  |  | Average |  | 7.9196 | 0.0015711 | 0.0012368 | 0.0004677 | 7.9183
Temperature readings [deg C] | 25 | 25.3 | 25.15 |  | 7.951 | 0.0025244 | 0.0011833 | 0.00047022 | 7.9498 | Real average voltage [V] | 0.0079183
Humidity readings [%] | 58.3 | 56.8 | 57.55 |  |  |  |  | Average [V] | 7.9341 | Real average current [A] | 7.9226e-06
|  |  |  | ORW-522.RA | [5 empty] | Real average current [mA] | 0.0079226
|  |  |  | Femto (ORE-546.RA) gain: 10^3; DC; L; FBW [Ohm] | Light [V] | SDEV [V] | Dark [V] | SDEV [V] | Nett [V]
|  |  |  | 999.46 | 1.8985 | 0.00046947 | 0.0012059 | 0.00068934 | 1.8973
[5 empty] | 1.9239 | 0.0018983 | 0.0024557 | 1.51e-05 | 1.9215
[8 empty] | Average [V] | 1.9094 | Real average voltage [V] | 0.0019094
[8 empty] | Current [A] | 0.0019104 | Real average current [A] | 1.9104e-06
[8 empty] | Current [mA] | 1.9104 | Real average current [mA] | 0.0019104
|  |  |  | ORW-522.RA Power responsivity at 905 nm [A/W] | 0.507 |  |  | Power | 0.0037681 | LIA measured irradiance [W.m-2] | 665.894 | 2.1323e-08 | 3.7681e-06 | Real irradiance [W.m-2] |  | 0.66589
[10 empty] | ORW-522.RA Irradiance responsivity at 905 nm [A/W.m-2] | 2.869e-06 |  | 0.66589
[10 empty] | ORW-522.RA Irradiance responsivity at 905 nm [V/W.m-2] | 0.0028674 | #VALUE!
Aperture area [m^2] |  | 5.6587e-06 | [7 empty] | Irradiance responsivity of HCA [A/W.m-2] | 1.1921e-05
[10 empty] | Irradiance responsivity of HCA [V/W.m-2] | 0.011915 | 0.011915
[10 empty] | Irradiance responsivity of HCA [mV/W.m-2] | 11.9149 | 11.9149
[10 empty] | Irradiance responsivity of HCA, with non-linearity correction [mV/W.m-2] | 11.8899 |  |  |  | 1
Calibration of Saab UUT against HCA-S-200M-SI (HCA)
Date of calibration | 17 - 21 Jan 2019
Used cryogenic radiometer's X-Y stage to position detector units
Used HeNe laser to define optical axis for detector units - to ensure exposure to same irradiance field
Temperature and humidity device (Hygroloc NMI00013647)
|  |  | Temperature readings (delta T) [deg C]: |  | 1 | varied between 24,5 deg C and 25,5 deg C
|  |  | Humidity readings [%]: |  |  | varied between 57,5 % and 58,6 %
Source | SAAB's pulsed source: EG&G InGaAs pulsed laser diode: PGAS3S09; wavelength = 905 nm +- 10 nm; wavelength temperature coefficient = 0,25 nm/deg C |  |  |  |  | With a temperature variation of 1 deg C over the calibration period, the wavelength variation is |  |  | 0,25 nm
Source operation: | Pulsed at 10 Hz using TTL provided by signal generator (Agilent 33220A) and power supplied with Agilent E3648A at 8 Volt DC
4700
HCA
Oscilloscope (ORE-096.RA) settings: |  |  |  | (367,65; 23,25) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [mV] | Nett [mV]
Bandwidth | 2GHz |  |  |  | 175.7552 | 1.142 |  |  | 175.7552
Sample rate | 1GS/s |  |  |  | 175.0555 | 0.6098 |  |  | 175.0555
Average of 200 samples |  |  |  |  | 174.9925 | 0.08669 |  |  | 174.9925 | HCA Gain[Ohm] | 20000 | Supplier specs
Termination | 50 Ohm | [7 empty] | 0 | Average HCA photovoltage [mV] | 176.5544
Coupling | DC | [6 empty] | Average [mV] | 175.4053 | Photocurrent [mA] | 8.8277e-06
[10 empty] | Photocurrent ratio (calibrated versus usage value) | 0.0011142
[11 empty] | Use correction factor used in Met 40 (2003)S35-S38: |  |  |  |  | 0.0007
SAAB UUT 4003
|  |  |  | (67,75; 9,75) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [mV] | Nett [mV]
[5 empty] | 539.4573 | 0.1118 |  |  | 539.4573
[5 empty] | 539.4395 | 0.7517 |  |  | 539.4395
[5 empty] | 537.6987 | 0.8873 |  |  | 537.6987
0
[8 empty] | Average [mV] | 538.8652
Non-lin. Corr.
|  |  |  | Irradiance responsivity of SAAB UUT 4003 [mV/W.m^2] | 36.604 | 36.5273 |  | Irradiance level [W.m^2] | 14.7215 | Percentage difference between corrected and not-corrected for non-linearity | 0.20956 | This calculation is based on Met. 37 (2000) 279-284 and Met. 40 (2003) S35-S38
|  |  |  | HCA | [6 empty] | Verification
|  |  |  | (367,65; 23,25) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [uV] | Nett [mV] |  | HCA
[5 empty] | 177.6613 | 0.6512 |  |  | 177.6613 |  | (367,65; 23,25) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [mV] | Nett [mV]
[5 empty] | 176.6264 | 0.7006 |  |  | 176.6264 |  |  | 175.4063 | 1.62 |  |  | 175.4063
[5 empty] | 176.0443 | 0.7591 |  |  | 176.0443 | [5 empty] | Average [mV] | 175.4063
[5 empty] | 175.4063 | 1.62 |  |  | 175.4063
[8 empty] | Average [mV] | 176.4346 |  | SAAB UUT 4003
[11 empty] | (67,75; 9,75) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [mV] | Nett [mV]
[12 empty] | 543.4816 | 1.161 |  |  | 543.4816
[15 empty] | Average [mV] | 543.4816
SAAB UUT 4020
|  |  |  | (262,8; 11,6) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [uV] | Nett [mV] |  | Irradiance responsivity of SAAB UUT 4003 [mV/W.m^2] | 36.9174
[5 empty] | 493.6013 | 1.374 |  |  | 493.6013
[5 empty] | 497.8624 | 0.8643 |  |  | 497.8624
[5 empty] | 495.4425 | 1.691 |  |  | 495.4425
[5 empty] | 495.8319 | 1.005 |  |  | 495.8319
[8 empty] | Average [mV] | 495.6845
|  |  |  | Irradiance responsivity of SAAB UUT 4020 [mV/W.m^2] | 33.4744 |  |  | Irradiance level [W.m^2] | 14.8079
HCA
|  |  |  | (367,85; 22,75) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [uV] | Nett [mV]
[5 empty] | 179.2821 | 1.364 |  |  | 179.2821 |  | Verification
[5 empty] | 177.058 | 1.915 |  |  | 177.058 |  | HCA
[5 empty] | 176.1982 | 1.054 |  |  | 176.1982 |  | (367,85; 22,75) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [mV] | Nett [mV]
[5 empty] | 175.3341 | 0.9169 |  |  | 175.3341 |  |  | 177.058 | 1.915 |  |  | 177.058
[8 empty] | Average [mV] | 176.9681 | [5 empty] | Average [mV] | 177.058
SAAB UUT 4003
[11 empty] | (67,75; 9,5) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [mV] | Nett [mV]
|  |  |  | SAAB UUT 4022 | [7 empty] | 543.8724 | 1.496 |  |  | 543.8724
|  |  |  | (262,8; 11,1) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [uV] | Nett [mV] | [5 empty] | Average [mV] | 543.8724
[5 empty] | 428.8345 | 1.068 |  |  | 428.8345
[5 empty] | 431.3482 | 0.823 |  |  | 431.3482 |  | Irradiance responsivity of SAAB UUT 4003 [mV/W.m^2] | 36.5993
[5 empty] | 429.132 | 1.228 |  |  | 429.132
[5 empty] | 419.6002 | 1.449 |  |  | 419.6002
[8 empty] | Average [mV] | 427.2287
|  |  |  | Irradiance responsivity of SAAB UUT 4022 [mV/W.m^2] | 28.7645 |  |  | Irradiance level [W.m^2] | 14.8527
HCA
|  |  |  | (367,65; 23,25) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [uV] | Nett [mV]
[5 empty] | 177.8278 | 0.9365 |  |  | 177.8278 |  | Verification
[5 empty] | 177.3314 | 1.006 |  |  | 177.3314 |  | HCA
[5 empty] | 175.6491 | 0.8882 |  |  | 175.6491 |  | (367,65; 23,25) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [mV] | Nett [mV]
[5 empty] | 176.142 | 0.6514 |  |  | 176.142 |  |  | 177.3314 | 1.006 |  |  | 177.3314
[8 empty] | Average [mV] | 176.7376 | [5 empty] | Average [mV] | 177.3314
SAAB UUT 4003
[11 empty] | (67,75; 9,7) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [mV] | Nett [mV]
|  |  |  | SAAB UUT 4030 | [7 empty] | 545.1959 | 1.439 |  |  | 545.1959
|  |  |  | (262,6; 11,45) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [uV] | Nett [mV] | [5 empty] | Average [mV] | 545.1959
[5 empty] | 503.0369 | 1.92 |  |  | 503.0369
[5 empty] | 504.4801 | 1.087 |  |  | 504.4801 |  | Irradiance responsivity of SAAB UUT 4003 [mV/W.m^2] | 36.6318
[5 empty] | 503.7673 | 1.316 |  |  | 503.7673
[5 empty] | 503.0206 | 1.169 |  |  | 503.0206
[8 empty] | Average [mV] | 503.5762
|  |  |  | Irradiance responsivity of SAAB UUT 4030 [mV/W.m^2] | 33.949 |  |  | Irradiance level [W.m^2] | 14.8333
HCA
|  |  |  | (367,85; 23,4) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [uV] | Nett [mV]
[5 empty] | 178.7048 | 0.9404 |  |  | 178.7048
[5 empty] | 176.8295 | 1.218 |  |  | 176.8295 |  | Verification
[5 empty] | 176.3488 | 0.7141 |  |  | 176.3488 |  | HCA
[5 empty] | 175.3926 | 1.527 |  |  | 175.3926 |  | (367,85; 23,4) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [mV] | Nett [mV]
[8 empty] | Average [mV] | 176.8189 |  |  | 176.8295 | 1.218 |  |  | 176.8295
[15 empty] | Average [mV] | 176.8295
SAAB UUT 4003
|  |  |  | SAAB UUT 4008 | [6 empty] | (67,85; 9,9) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [mV] | Nett [mV]
|  |  |  | (255,5; 15,65) | Light [mV] | SDEV [mV] | Dark [mV] | SDEV [uV] | Nett [mV] |  |  | 545.8986 | 1.005 |  |  | 545.8986
[5 empty] | 541.5897 | 1.788 |  |  | 541.5897 | [5 empty] | Average [mV] | 545.8986
[5 empty] | 532.8158 | 1.821 |  |  | 532.8158
[5 empty] | 520.9149 | 0.9729 |  |  | 520.9149 |  | Irradiance responsivity of SAAB UUT 4003 [mV/W.m^2] | 36.7831
[5 empty] | 521.1667 | 1.384 |  |  | 521.1667
[8 empty] | Average [mV] | 529.1218
|  |  |  | Irradiance responsivity of SAAB UUT 4008 [mV/W.m^2] | 35.6548 |  |  | Irradiance level [W.m^2] | 14.8401
Irradiance responsivity of photodiode modules [mV/W.m^2]
[10 empty] | Unit 4003 | 36.604
[10 empty] | Unit 4008 | 35.6548
[10 empty] | Unit 4020 | 33.4744
[10 empty] | Unit 4022 | 28.7645
[10 empty] | Unit 4030 | 33.949
[10 empty] | Unit 4031 | NA
[10 empty] | Average irradiance level used during calibration [W.m^2] | 14.8111

## Averaged data
Wavelength [nm] | Data sets [A/W] | [6 empty] | Comparison with previous certifiates
| Set 1 | Set 2 | Set 3 | Set 4 | Average of sets [A/W] | Pyro absorbance (normalised) [1] | Response corrected for absorbance. OR\SR-5416 [A/W] | OR\RA-2899 [A/W] | Uc [%] | Difference between OR\RA-2899 & OR\SR-5416 [A/W] | OR\SR-3592 [A/W] | Uc [%] | Uc [A/W] | Difference between OR\SR-3592 & OR\SR-5416 [A/W]
2.6%
600 | 0.33253 | 0.3328 | 0.3332 |  | 0.33284 | 1.0001 | 0.33287 | 0.332 | 0.008632 | 0.00086772 | 0.321 | 2% | 0.00642 | 0.011868
610 | 0.3371 | 0.33448 | 0.33677 |  | 0.33612 | 1.0001 | 0.33614 | 0.337 | 0.008762 | 0.00086151 | 0.325 | 2% | 0.0065 | 0.011138
620 | 0.34397 | 0.34849 | 0.34529 |  | 0.34592 | 1.0001 | 0.34593 | 0.342 | 0.008892 | 0.003933 | 0.33 | 2% | 0.0066 | 0.015933
630 | 0.35024 | 0.34708 | 0.35275 |  | 0.35002 | 1 | 0.35003 | 0.347 | 0.009022 | 0.0030283 | 0.334 | 2% | 0.00668 | 0.016028
640 | 0.35625 | 0.35066 | 0.36167 |  | 0.35619 | 0.99996 | 0.35618 | 0.354 | 0.009204 | 0.0021763 | 0.339 | 2% | 0.00678 | 0.017176
650 | 0.36054 | 0.35931 | 0.36011 |  | 0.35999 | 0.99988 | 0.35994 | 0.36 | 0.00936 | 5.6202e-05 | 0.345 | 2% | 0.0069 | 0.014944
660 | 0.3646 | 0.36265 | 0.36685 |  | 0.3647 | 0.99978 | 0.36462 | 0.366 | 0.009516 | 0.0013794 | 0.349 | 2% | 0.00698 | 0.015621
670 | 0.37437 | 0.36953 | 0.37258 |  | 0.37216 | 0.99964 | 0.37203 | 0.372 | 0.009672 | 2.5438e-05 | 0.353 | 2% | 0.00706 | 0.019025
680 | 0.37852 | 0.37778 | 0.37859 |  | 0.3783 | 0.99948 | 0.3781 | 0.378 | 0.009828 | 0.0001018 | 0.357 | 2% | 0.00714 | 0.021102
690 | 0.38839 | 0.38108 | 0.38453 |  | 0.38467 | 0.99928 | 0.38439 | 0.385 | 0.01001 | 0.00061045 | 0.364 | 2% | 0.00728 | 0.02039
700 | 0.3898 | 0.38862 | 0.39017 |  | 0.38953 | 0.99904 | 0.38916 | 0.391 | 0.010166 | 0.0018435 | 0.37 | 2% | 0.0074 | 0.019157
710 | 0.39617 | 0.39937 | 0.39524 |  | 0.39693 | 0.99875 | 0.39643 | 0.398 | 0.010348 | 0.0015677 | 0.374 | 2% | 0.00748 | 0.022432
720 | 0.40323 | 0.40855 | 0.40191 |  | 0.40456 | 0.99842 | 0.40393 | 0.404 | 0.010504 | 7.3607e-05 | 0.38 | 2% | 0.0076 | 0.023926
730 | 0.40629 | 0.40312 | 0.4081 |  | 0.40584 | 0.99805 | 0.40505 | 0.41 | 0.01066 | 0.0049543 | 0.384 | 2% | 0.00768 | 0.021046
740 | 0.41489 | 0.4148 | 0.41426 |  | 0.41465 | 0.99765 | 0.41368 | 0.417 | 0.010842 | 0.0033243 | 0.389 | 2% | 0.00778 | 0.024676
750 | 0.42128 | 0.42757 | 0.4196 |  | 0.42282 | 0.99722 | 0.42164 | 0.421 | 0.010946 | 0.00064154 | 0.393 | 2% | 0.00786 | 0.028642
760 | 0.42352 | 0.42271 | 0.42692 |  | 0.42439 | 0.99677 | 0.42302 | 0.427 | 0.011102 | 0.0039848 | 0.399 | 2% | 0.00798 | 0.024015
770 | 0.43064 | 0.42849 | 0.43355 |  | 0.43089 | 0.99631 | 0.4293 | 0.432 | 0.011232 | 0.0026979 | 0.403 | 2% | 0.00806 | 0.026302
780 | 0.43536 | 0.43989 | 0.44067 |  | 0.43864 | 0.99583 | 0.43681 | 0.439 | 0.011414 | 0.0021867 | 0.408 | 2% | 0.00816 | 0.028813
790 | 0.44956 | 0.43964 | 0.4482 |  | 0.4458 | 0.99536 | 0.44373 | 0.445 | 0.01157 | 0.0012678 | 0.413 | 2% | 0.00826 | 0.030732
800 | 0.45198 | 0.45049 | 0.44863 |  | 0.45037 | 0.99489 | 0.44806 | 0.45 | 0.0117 | 0.0019352 | 0.418 | 2% | 0.00836 | 0.030065
810 | 0.45514 | 0.44717 | 0.45549 |  | 0.4526 | 0.9944 | 0.45007 | 0.456 | 0.011856 | 0.0059341 | 0.423 | 2% | 0.00846 | 0.027066
820 | 0.46528 | 0.46271 | 0.46279 |  | 0.46359 | 0.99388 | 0.46076 | 0.461 | 0.011986 | 0.00024136 | 0.428 | 2% | 0.00856 | 0.032759
830 | 0.47006 | 0.47004 | 0.46853 |  | 0.46954 | 0.99334 | 0.46642 | 0.468 | 0.012168 | 0.0015831 | 0.434 | 2% | 0.00868 | 0.032417
840 | 0.473 | 0.4745 | 0.47492 |  | 0.47414 | 0.99279 | 0.47072 | 0.474 | 0.012324 | 0.0032807 | 0.439 | 2% | 0.00878 | 0.031719
850 | 0.47892 | 0.4787 | 0.4781 |  | 0.47857 | 0.99223 | 0.47485 | 0.479 | 0.012454 | 0.004145 | 0.446 | 2% | 0.00892 | 0.028855
860 | 0.48493 | 0.48302 | 0.48498 |  | 0.48431 | 0.99168 | 0.48028 | 0.486 | 0.012636 | 0.0057245 | 0.451 | 2% | 0.00902 | 0.029276
870 | 0.48951 | 0.49898 | 0.49201 |  | 0.4935 | 0.99114 | 0.48913 | 0.491 | 0.012766 | 0.0018741 | 0.456 | 2% | 0.00912 | 0.033126
880 | 0.49736 | 0.49562 | 0.49919 |  | 0.49739 | 0.99062 | 0.49272 | 0.496 | 0.012896 | 0.0032758 | 0.461 | 2% | 0.00922 | 0.031724
890 | 0.50322 | 0.50043 | 0.50714 |  | 0.5036 | 0.99014 | 0.49863 | 0.501 | 0.013026 | 0.0023662 | 0.467 | 2% | 0.00934 | 0.031634
900 | 0.51296 | 0.50689 | 0.5106 |  | 0.51015 | 0.9897 | 0.50489 | 0.507 | 0.013182 | 0.0021051 | 0.472 | 2% | 0.00944 | 0.032895
910 | 0.51593 | 0.51837 | 0.51652 |  | 0.51694 | 0.9893 | 0.51141 | 0.513 | 0.013338 | 0.0015935 | 0.477 | 2% | 0.00954 | 0.034406
920 | 0.52074 | 0.52363 | 0.52385 |  | 0.52274 | 0.98893 | 0.51695 | 0.52 | 0.01352 | 0.0030451 | 0.482 | 2% | 0.00964 | 0.034955
930 | 0.52887 | 0.52948 | 0.52882 |  | 0.52906 | 0.98859 | 0.52302 | 0.525 | 0.01365 | 0.0019791 | 0.488 | 2% | 0.00976 | 0.035021
940 | 0.5335 | 0.53614 | 0.53585 |  | 0.53516 | 0.98827 | 0.52889 | 0.531 | 0.013806 | 0.0021131 | 0.492 | 2% | 0.00984 | 0.036887
950 | 0.53872 | 0.54048 | 0.54055 | 0.54185 | 0.5404 | 0.98797 | 0.5339 | 0.535 | 0.01391 | 0.0011003 | 0.496 | 2% | 0.00992 | 0.0379
960 | 0.5414 | 0.54504 | 0.5461 | 0.54728 | 0.54495 | 0.98768 | 0.53824 | 0.535 | 0.01391 | 0.0032423 | 0.504 | 3% | 0.01512 | 0.034242
970 | 0.54336 | 0.55479 | 0.54868 | 0.55108 | 0.54948 | 0.98741 | 0.54256 | 0.536 | 0.013936 | 0.0065552 | 0.507 | 3% | 0.01521 | 0.035555
980 | 0.54366 | 0.55405 | 0.54942 | 0.5527 | 0.54996 | 0.98713 | 0.54288 | 0.535 | 0.01391 | 0.0078814 | 0.505 | 3% | 0.01515 | 0.037881
990 | 0.54102 | 0.54822 | 0.54554 | 0.54975 | 0.54613 | 0.98686 | 0.53896 | 0.531 | 0.013806 | 0.0079584 | 0.499 | 3% | 0.01497 | 0.039958
1000 | 0.53351 | 0.54 | 0.53721 | 0.54037 | 0.53777 | 0.98659 | 0.53056 | 0.528 | 0.013728 | 0.0025623 | 0.487 | 3% | 0.01461 | 0.043562
1010 | 0.51646 |  | 0.52264 | 0.52463 | 0.52124 | 0.98632 | 0.51411 |  |  |  | 0.467 | 3% | 0.01401 | 0.04711
1020 | 0.48931 |  | 0.49556 | 0.49446 | 0.49311 | 0.98605 | 0.48623 |  |  |  | 0.438 | 3% | 0.01314 | 0.048229
1030 | 0.45176 |  | 0.4574 | 0.45541 | 0.45486 | 0.98579 | 0.44839 |  |  |  | 0.397 | 3% | 0.01191 | 0.051393
1040 | 0.39856 |  | 0.40367 | 0.4044 | 0.40221 | 0.98554 | 0.39639 |  |  |  | 0.348 | 3% | 0.01044 | 0.048393
1050 | 0.34235 |  | 0.3456 | 0.34573 | 0.34456 | 0.98529 | 0.33949 |  |  |  | 0.291 | 4% | 0.01164 | 0.048493
1060 | 0.27967 |  | 0.28292 | 0.28336 | 0.28198 | 0.98505 | 0.27777 |  |  |  | 0.232 | 4% | 0.00928 | 0.045768
1070 | 0.22892 |  | 0.23147 | 0.2309 | 0.23043 | 0.98482 | 0.22693 |  |  |  | 0.189 | 4% | 0.00756 | 0.037931
1080 | 0.1887 |  | 0.19335 | 0.19296 | 0.19167 | 0.9846 | 0.18872 |  |  |  | 0.156 | 4% | 0.00624 | 0.03272
1090 | 0.19035 |  | 0.1591 | 0.15891 | 0.15901 | 0.98438 | 0.15652 |  |  |  | 0.126 | 4% | 0.00504 | 0.030521
1100 |  |  | 0.12867 | 0.12868 | 0.12867 | 0.98417 | 0.12664 |  |  |  | 0.102 | 4% | 0.00408 | 0.024637
633 |  |  |  |  | 0.35055 | 1 | 0.35055
905 | 0.50953 | 0.51491 | 0.51373 |  | 0.51272 | 0.9895 | 0.50734

## 13 Dec
Detectors positions | x | y
STD: UPD-50-UD | 342 | 24
UUT 4020 | 262.25 | 11.5
UUT 4003 | 67.4 | 9.4